# LPA3-NAVAIR Online Calibration Workflow

- Optional preflush units 2 mins on rig 303

For each unit perform electrical tests, these could be done offline:

- File new / LPA3 should appear in com ports list

- If won't connect use MPROG to set lpa3 USB ID

- Check power button+light works, turn off power

- Firmware update 5.17, check

- Check takes power, charges @1.8A from 19VDC Lab PSU via Netwarrior adapter cable

- LPAView LPA3 Calibration - change model to LPA3-Navair, APPLY, allow to reset

- INITIALIZE PRODUCT SETTINGS, set S/N e.g. 63010046

- BOTTLE SAMPLE MODE (view add reference control panel)

- ISO reporting format on LPA – user settings change nas to iso, bottle sample)

- Enable thresholds.

- Pressurize & flush clean oil

- Noise check – generally around 30. Sets LED level at calibration, noise threshold in cal certificate. Should be < 40. |LED level should be >8000.

- Load reference distribution for Partistan 37d

# Offline Calibration

Partistan prep

- Shake 1 minute

- Sonicate 1 minute

- Turn 1 minute

- Sonicate 1 minute

- Fill 4x ultraclean bottles, ~25ml at a time.

Per bottle

- Shake 1min

- Vacuum 1m20

- Pressurize until cuts out

TWO bottles Reference Fluid per calibration

- Bottle 1

- Ref control panel / apply edit – check target counts in below box

- Check bottle sampling in user settings

- START test on Partistan

- When test completed, AUTO ADJUST button

- Thresholds will auto adjust

- Recalc on LPA3, PRINT

- Bottle 2

- START Test

- Verification test completes

- Add >38 thresholds

- CREATE cal cert

TBC