## Clause 6
Clause 6 - Sizing Calibration - Initial Threshold Setting
|  |  | Reference Fluid Dilution Calculator | [11 empty] | Diluted | 89% | Captured
RM | Partistan 34D Lot 34D |  |  | Target Concentration | 0.3 | [9 empty] | 0.3 | 50 | ml
Date | 2025-11-27 |  |  | Fluid Density | 0.868 | [5 empty] | um(c) | SRM2806D | 34D | RM USED | Diluted | Counted
Operator | JMD | [9 empty] | 1.5 | 67392.07 |  | 67392.07 | 20217.621 | 899684.1345
APC Model | REF |  |  | Bottle Number | B0 | B1 | B2 | B3 |  |  | 2 | 41165 |  | 41165 | 12349.5 | 549552.75
APC S/N | 11171-00001 |  |  | Ref Fluid to Add / g | 26.04 | 26.04 | 26.04 | 26.04 |  |  | 3 | 17975.47 | 17634.1 | 17975.47 | 5392.641 | 239972.5245
Sensor Model | REF |  |  | Ref Fluid Added / g | 26.08 | 26.08 | 26.11 | 26.16 |  |  | 4 | 10898.56 | 10750.3 | 10898.56 | 3269.568 | 145495.776
Sensor S/N | 11171-00001 |  |  | Diluent to add / g | 60.8533 | 60.8533 | 60.9233 | 61.04 |  |  | 5 | 6986.32 | 7121.1 | 6986.32 | 2095.896 | 93267.372
Bottle Sampler | BS110-HP D2 BS2044 |  |  | Diluent Added / g | 61 | 61.18 | 61.2 | 61.09 |  |  | 6 | 4679.65 | 4729.1 | 4679.65 | 1403.895 | 62473.3275
|  |  |  | Final Concentration | 0.29949 | 0.29888 | 0.29905 | 0.29983 |  |  | 7 | 3163.23 | 3199.3 | 3163.23 | 948.969 | 42229.1205
Threshold Noise | 110 | [9 empty] | 8 | 2175.47 | 2280.6 | 2175.47 | 652.641 | 29042.5245
[11 empty] | 9 | 1512.11 | 1615 | 1512.11 | 453.633 | 20186.6685
Lowest threshold J | 165 | [9 empty] | 10 | 1079.78 | 1130.5 | 1079.78 | 323.934 | 14415.063
Highest threshold H | 14000 | [9 empty] | 11 | 786.62 | 816.2 | 786.62 | 235.986 | 10501.377
Threshold Factor K | 1.1655 | [9 empty] | 12 | 586.06 | 618.6 | 586.06 | 175.818 | 7823.901
[11 empty] | 13 | 449.93 | 481.9 | 449.93 | 134.979 | 6006.5655
Threshold# | Initial Value | INT | Reversed | [7 empty] | 14 | 352.62 | 384.7 | 352.62 | 105.786 | 4707.477
4 | 165 | 165 | 14000 | 1.5x noise | [6 empty] | 15 | 282.04 | 314.6 | 282.04 | 84.612 | 3765.234
5 | 192.3043 | 192 | 12012.2142 | [7 empty] | 16 | 229.42 | 261.8 | 229.42 | 68.826 | 3062.757
6 | 224.1268 | 224 | 10306.6635 | [7 empty] | 17 | 186.1 | 221.2 | 186.1 | 55.83 | 2484.435
7 | 261.2154 | 261 | 8843.275 | [7 empty] | 18 | 153.87 | 188.7 | 153.87 | 46.161 | 2054.1645
8 | 304.4415 | 304 | 7587.6652 | [7 empty] | 19 | 128.62 | 161 | 128.62 | 38.586 | 1717.077
9 | 354.8206 | 355 | 6510.3328 | [7 empty] | 20 | 107.35 | 136.4 | 107.35 | 32.205 | 1433.1225
10 | 413.5364 | 414 | 5585.9652 | [7 empty] | 21 | 88.64 | 114.4 | 88.64 | 26.592 | 1183.344
11 | 481.9686 | 482 | 4792.8436 | [7 empty] | 22 | 75.56 | 95.3 | 75.56 | 22.668 | 1008.726
12 | 561.7249 | 562 | 4112.3331 | [7 empty] | 23 | 61.94 | 79 | 61.94 | 18.582 | 826.899
13 | 654.6794 | 655 | 3528.4447 | [7 empty] | 24 | 50.86 | 65.3 | 50.86 | 15.258 | 678.981
14 | 763.016 | 763 | 3027.4595 | [7 empty] | 25 | 41.61 | 53.8 | 41.61 | 12.483 | 555.4935
15 | 889.2802 | 889 | 2597.6066 | [7 empty] | 26 | 33.44 | 44.6 | 33.44 | 10.032 | 446.424
16 | 1036.4386 | 1036 | 2228.7862 | [7 empty] | 27 | 26.08 | 37 | 26.08 | 7.824 | 348.168
17 | 1207.9488 | 1208 | 1912.3327 | [7 empty] | 28 | 20.41 | 30.7 | 20.41 | 6.123 | 272.4735
18 | 1407.8407 | 1408 | 1640.8107 | [7 empty] | 29 | 16.39 | 25.2 | 16.39 | 4.917 | 218.8065
19 | 1640.8107 | 1641 | 1407.8407 | [7 empty] | 30 | 13.94 | 20.1 | 13.94 | 4.182 | 186.099
20 | 1912.3327 | 1912 | 1207.9488
21 | 2228.7862 | 2229 | 1036.4386
22 | 2597.6066 | 2598 | 889.2802
23 | 3027.4595 | 3027 | 763.016
24 | 3528.4447 | 3528 | 654.6794
25 | 4112.3331 | 4112 | 561.7249
26 | 4792.8436 | 4793 | 481.9686
27 | 5585.9652 | 5586 | 413.5364
28 | 6510.3328 | 6510 | 354.8206
29 | 7587.6652 | 7588 | 304.4415
30 | 8843.275 | 8843 | 261.2154
38 | 10306.6635 | 10307 | 224.1268
50 | 12012.2142 | 12012 | 192.3043
70 | 14000 | 14000 | 165 | ~30um
INITIAL Thresholds to Load into Instrument
Number | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 38 | 50 | 70
Value | 165 | 192 | 224 | 261 | 304 | 355 | 414 | 482 | 562 | 655 | 763 | 889 | 1036 | 1208 | 1408 | 1641 | 1912 | 2229 | 2598 | 3027 | 3528 | 4112 | 4793 | 5586 | 6510 | 7588 | 8843 | 10307 | 12012 | 14000
TURN ON RAW COUNTS AND COINCIDENCE CORRECTION FOR THRESHOLD CALIBRATION

## 6 - B1
Data Preprocessing and Quality Assessment per ISO 11171 Clause 6.6
BOTTLE 1 | [6 empty] | CHECK CALIBRATION THRESHOLDS ENTERED ON PARTICLE COUNTER
Serial # | 11171-00009 |  | Dilution Ratio | 0.29888
Date | 2025-11-28
RM | 34D
BOTTLE | 1
| Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | X / ml | NTOTAL | CV | DQMAX | DQ | X̅N | NC
Scanned ml | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 43.1227
count_4 | 4082.3649 | 4059.984 | 4137.4312 | 4038.5768 | 4238.2025 | 4111.3119 | 177291.0417 | 1.739% | 11% | 4.856% | 13755.8694 | 5
count_5 | 3566.3496 | 3548.7479 | 3605.0391 | 3518.5205 | 3680.9667 | 3583.9247 | 154548.6622 | 1.563% | 11% | 4.533% | 11991.3065 | 5
count_6 | 3102.3704 | 3091.5916 | 3126.1745 | 3046.8468 | 3206.2843 | 3114.6535 | 134312.4002 | 1.687% | 11% | 5.119% | 10421.1912 | 5
count_7 | 2692.9628 | 2685.4213 | 2694.0313 | 2648.6643 | 2779.2995 | 2700.0758 | 116434.6741 | 1.591% | 11% | 4.838% | 9034.0728 | 5
count_8 | 2317.602 | 2304.6868 | 2314.262 | 2272.6095 | 2402.5348 | 2322.339 | 100145.6267 | 1.857% | 11% | 5.595% | 7770.2187 | 5
count_9 | 1977.737 | 1973.5635 | 1988.4819 | 1947.1157 | 2045.0324 | 1986.3861 | 85658.4153 | 1.627% | 11% | 4.929% | 6646.1676 | 5
count_10 | 1676.2652 | 1677.6149 | 1685.3125 | 1653.6095 | 1737.6355 | 1686.0875 | 72708.717 | 1.652% | 11% | 4.983% | 5641.4109 | 5
count_11 | 1410.2891 | 1415.5807 | 1409.0223 | 1387.6194 | 1458.3297 | 1416.1682 | 61069.0576 | 1.634% | 11% | 4.993% | 4738.2991 | 5
count_12 | 1168.339 | 1165.3477 | 1164.5718 | 1151.4387 | 1208.3761 | 1171.6147 | 50523.237 | 1.645% | 11% | 4.86% | 3920.0574 | 5
count_13 | 952.4673 | 953.9559 | 956.9216 | 948.6213 | 982.5007 | 958.8934 | 41350.1106 | 1.262% | 11% | 3.533% | 3208.3219 | 5
count_14 | 767.9864 | 766.3957 | 774.3047 | 767.7022 | 788.0415 | 772.8861 | 33328.968 | 1.043% | 11% | 2.801% | 2585.9678 | 5
count_15 | 618.7597 | 611.7188 | 616.0291 | 613.2675 | 615.2525 | 615.0055 | 26520.7237 | 0.3918% | 11% | 1.145% | 2057.7216 | 5
count_16 | 484.7454 | 475.4885 | 479.326 | 486.349 | 479.2686 | 481.0355 | 20743.5693 | 0.8253% | 11% | 2.258% | 1609.4769 | 5
count_17 | 373.9119 | 372.4851 | 370.0789 | 377.66 | 369.0827 | 372.6437 | 16069.4186 | 0.8144% | 11% | 2.302% | 1246.8133 | 5
count_18 | 290.0021 | 290.2198 | 288.8646 | 292.8183 | 286.3 | 289.641 | 12490.112 | 0.7296% | 11% | 2.25% | 969.0978 | 5
count_19 | 221.7876 | 229.4948 | 223.8008 | 226.7794 | 220.8305 | 224.5386 | 9682.7212 | 1.429% | 11.3% | 3.859% | 751.2746 | 5
count_20 | 171.4417 | 177.4775 | 176.5027 | 174.2693 | 171.8717 | 174.3126 | 7516.8367 | 1.382% | 11.3% | 3.463% | 583.2253 | 5
count_21 | 137.5156 | 140.0113 | 138.8949 | 139.0715 | 135.6399 | 138.2266 | 5960.7119 | 1.099% | 11.3% | 3.162% | 462.4868 | 5
count_22 | 106.487 | 110.1071 | 108.2089 | 107.5425 | 103.7651 | 107.2221 | 4623.7117 | 1.951% | 11.9% | 5.915% | 358.7501 | 5
count_23 | 81.4952 | 86.5045 | 83.1755 | 83.1219 | 78.7698 | 82.6134 | 3562.5145 | 3.05% | 11.9% | 9.363% | 276.4126 | 5
count_24 | 65.9205 | 68.287 | 66.6788 | 67.0708 | 60.7685 | 65.7451 | 2835.1101 | 3.96% | 11.9% | 11.44% | 219.9739 | 5
count_25 | 48.7764 | 52.0173 | 49.7207 | 51.2489 | 44.9458 | 49.3418 | 2127.7543 | 5.012% | 11.9% | 14.33% | 165.0907 | 5
count_26 | 36.8237 | 40.3306 | 38.4153 | 38.6373 | 32.3334 | 37.3081 | 1608.8268 | 7.302% | 13.4% | 21.44% | 124.8276 | 5
count_27 | 28.6139 | 32.5395 | 29.9939 | 29.9239 | 23.2755 | 28.8693 | 1244.9242 | 10.64% | 13.4% | 32.09% | 96.5927 | 5
count_28 | 23.0601 | 23.7171 | 22.8415 | 22.7009 | 17.3133 | 21.9266 | 945.5346 | 10.64% | 15.6% | 29.21% | 73.3633 | 5
count_29 | 15.8161 | 18.1029 | 17.0735 | 16.739 | 13.3003 | 16.2064 | 698.8627 | 10.04% | 15.6% | 29.63% | 54.2242 | 5
count_30 | 10.6246 | 13.7491 | 13.0358 | 11.6944 | 9.1726 | 11.6553 | 502.6079 | 14.11% | 15.6% | 39.27% | 38.9969 | 5
count_38 | 7.9684 | 10.4264 | 9.8057 | 8.5988 | 6.1915 | 8.5982 | 370.7764 | 17.24% | 19.3% | 49.25% | 28.7682 | 5
count_50 | 5.7952 | 7.6766 | 7.3831 | 5.8472 | 3.8984 | 6.1201 | 263.915 | 22.08% | 19.3% | 61.73% | 20.477 | 5
count_70 | 4.2257 | 5.385 | 4.7298 | 4.3567 | 3.0958 | 4.3586 | 187.955 | 17.18% | 27.5% | 52.52% | 14.5833 | 5
Table C.2
| Mean counted X |  | CV_MAX
X >= | X < | DQMAX / SAMPLE | N=3 | N=6 | N=10 | N=20 | N=40
10000 |  | 11 | 7.8 | 5.3 | 4 | 2.7 | 1.8
5000 | 10000 | 11.3 | 7.8 | 5.3 | 4 | 2.7 | 1.9
2000 | 5000 | 11.9 | 7.9 | 5.4 | 4.1 | 2.9 | 2.1
1000 | 2000 | 13.4 | 8 | 5.6 | 4.3 | 3.1 | 2.4
500 | 1000 | 15.6 | 8.2 | 5.9 | 4.7 | 3.6 | 2.9
200 | 500 | 19.3 | 8.9 | 6.8 | 5.7 | 4.7 | 4.1
100 | 200 | 27.5 | 10 | 8 | 7 | 6.1 | 5.5
| Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5
NAME\mL | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216
4 | 33813 | 35435 | 35865 | 35225 | 36964 | 33813 | 35435 | 35865 | 35225 | 36964 | 33813 | 35435 | 35865 | 35225 | 36964
5 | 29539 | 30973 | 31250 | 30689 | 32104 | 29539 | 30973 | 31250 | 30689 | 32104 | 29539 | 30973 | 31250 | 30689 | 32104
6 | 25696 | 26983 | 27099 | 26575 | 27964 | 25696 | 26983 | 27099 | 26575 | 27964 | 25696 | 26983 | 27099 | 26575 | 27964
7 | 22305 | 23438 | 23353 | 23102 | 24240 | 22305 | 23438 | 23353 | 23102 | 24240 | 22305 | 23438 | 23353 | 23102 | 24240
8 | 19196 | 20115 | 20061 | 19822 | 20954 | 19196 | 20115 | 20061 | 19822 | 20954 | 19196 | 20115 | 20061 | 19822 | 20954
9 | 16381 | 17225 | 17237 | 16983 | 17836 | 16381 | 17225 | 17237 | 16983 | 17836 | 16381 | 17225 | 17237 | 16983 | 17836
10 | 13884 | 14642 | 14609 | 14423 | 15155 | 13884 | 14642 | 14609 | 14423 | 15155 | 13884 | 14642 | 14609 | 14423 | 15155
11 | 11681 | 12355 | 12214 | 12103 | 12719 | 11681 | 12355 | 12214 | 12103 | 12719 | 11681 | 12355 | 12214 | 12103 | 12719
12 | 9677 | 10171 | 10095 | 10043 | 10539 | 9677 | 10171 | 10095 | 10043 | 10539 | 9677 | 10171 | 10095 | 10043 | 10539
13 | 7889 | 8326 | 8295 | 8274 | 8569 | 7889 | 8326 | 8295 | 8274 | 8569 | 7889 | 8326 | 8295 | 8274 | 8569
14 | 6361 | 6689 | 6712 | 6696 | 6873 | 6361 | 6689 | 6712 | 6696 | 6873 | 6361 | 6689 | 6712 | 6696 | 6873
15 | 5125 | 5339 | 5340 | 5349 | 5366 | 5125 | 5339 | 5340 | 5349 | 5366 | 5125 | 5339 | 5340 | 5349 | 5366
16 | 4015 | 4150 | 4155 | 4242 | 4180 | 4015 | 4150 | 4155 | 4242 | 4180 | 4015 | 4150 | 4155 | 4242 | 4180
17 | 3097 | 3251 | 3208 | 3294 | 3219 | 3097 | 3251 | 3208 | 3294 | 3219 | 3097 | 3251 | 3208 | 3294 | 3219
18 | 2402 | 2533 | 2504 | 2554 | 2497 | 2402 | 2533 | 2504 | 2554 | 2497 | 2402 | 2533 | 2504 | 2554 | 2497
19 | 1837 | 2003 | 1940 | 1978 | 1926 | 1837 | 2003 | 1940 | 1978 | 1926 | 1837 | 2003 | 1940 | 1978 | 1926
20 | 1420 | 1549 | 1530 | 1520 | 1499 | 1420 | 1549 | 1530 | 1520 | 1499 | 1420 | 1549 | 1530 | 1520 | 1499
21 | 1139 | 1222 | 1204 | 1213 | 1183 | 1139 | 1222 | 1204 | 1213 | 1183 | 1139 | 1222 | 1204 | 1213 | 1183
22 | 882 | 961 | 938 | 938 | 905 | 882 | 961 | 938 | 938 | 905 | 882 | 961 | 938 | 938 | 905
23 | 675 | 755 | 721 | 725 | 687 | 675 | 755 | 721 | 725 | 687 | 675 | 755 | 721 | 725 | 687
24 | 546 | 596 | 578 | 585 | 530 | 546 | 596 | 578 | 585 | 530 | 546 | 596 | 578 | 585 | 530
25 | 404 | 454 | 431 | 447 | 392 | 404 | 454 | 431 | 447 | 392 | 404 | 454 | 431 | 447 | 392
26 | 305 | 352 | 333 | 337 | 282 | 305 | 352 | 333 | 337 | 282 | 305 | 352 | 333 | 337 | 282
27 | 237 | 284 | 260 | 261 | 203 | 237 | 284 | 260 | 261 | 203 | 237 | 284 | 260 | 261 | 203
28 | 191 | 207 | 198 | 198 | 151 | 191 | 207 | 198 | 198 | 151 | 191 | 207 | 198 | 198 | 151
29 | 131 | 158 | 148 | 146 | 116 | 131 | 158 | 148 | 146 | 116 | 131 | 158 | 148 | 146 | 116
30 | 88 | 120 | 113 | 102 | 80 | 88 | 120 | 113 | 102 | 80 | 88 | 120 | 113 | 102 | 80
38 | 66 | 91 | 85 | 75 | 54 | 66 | 91 | 85 | 75 | 54 | 66 | 91 | 85 | 75 | 54
50 | 48 | 67 | 64 | 51 | 34 | 48 | 67 | 64 | 51 | 34 | 48 | 67 | 64 | 51 | 34
70 | 35 | 47 | 41 | 38 | 27 | 35 | 47 | 41 | 38 | 27 | 35 | 47 | 41 | 38 | 27
8.2827 | 33813 | 29539 | 25696 | 22305 | 19196 | 16381 | 13884 | 11681 | 9677 | 7889 | 6361 | 5125 | 4015 | 3097 | 2402 | 1837 | 1420 | 1139 | 882 | 675 | 546 | 404 | 305 | 237 | 191 | 131 | 88 | 66 | 48 | 35
8.7279 | 35435 | 30973 | 26983 | 23438 | 20115 | 17225 | 14642 | 12355 | 10171 | 8326 | 6689 | 5339 | 4150 | 3251 | 2533 | 2003 | 1549 | 1222 | 961 | 755 | 596 | 454 | 352 | 284 | 207 | 158 | 120 | 91 | 67 | 47
8.6684 | 35865 | 31250 | 27099 | 23353 | 20061 | 17237 | 14609 | 12214 | 10095 | 8295 | 6712 | 5340 | 4155 | 3208 | 2504 | 1940 | 1530 | 1204 | 938 | 721 | 578 | 431 | 333 | 260 | 198 | 148 | 113 | 85 | 64 | 41
8.7221 | 35225 | 30689 | 26575 | 23102 | 19822 | 16983 | 14423 | 12103 | 10043 | 8274 | 6696 | 5349 | 4242 | 3294 | 2554 | 1978 | 1520 | 1213 | 938 | 725 | 585 | 447 | 337 | 261 | 198 | 146 | 102 | 75 | 51 | 38
8.7216 | 36964 | 32104 | 27964 | 24240 | 20954 | 17836 | 15155 | 12719 | 10539 | 8569 | 6873 | 5366 | 4180 | 3219 | 2497 | 1926 | 1499 | 1183 | 905 | 687 | 530 | 392 | 282 | 203 | 151 | 116 | 80 | 54 | 34 | 27
8.6736 | 33986 | 29686 | 25849 | 22427 | 19306 | 16497 | 13917 | 11670 | 9608 | 7881 | 6344 | 5001 | 3881 | 2987 | 2311 | 1766 | 1370 | 1058 | 811 | 647 | 489 | 367 | 271 | 190 | 131 | 99 | 74 | 59 | 35 | 22
Export into temportary CSV file 15 results from LPA-View, consisting of 3x bottles 5 portion samples. Exclude the portion 0 averages.
Autofilter, sort by ascending date.
Copy / special paste the data block into this page below, ensuring appears in correct slots
Bottle 2 and 3 pages should be linked to correct regions of pasted data block.
Data quality evalution tables should pick up correct regions of pasted data block
If needed, on each bottle page can delete one indivual portion result per size line in order tp meet data quality requirements

## 6 - B2
Data Preprocessing and Quality Assessment per ISO 11171 Clause 6.6
BOTTLE 2
Serial # | 11171-00009
Date | 2025-11-28
RM | 34D
| Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | X / ml | NTOTAL | CV | DQMAX | DQ | X̅N | NC
Scanned ml | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 43.1227
count_4 | 4082.3649 | 4059.984 | 4137.4312 | 4038.5768 | 4238.2025 | 4111.3119 | 177291.0417 | 1.739% | 11% | 4.856% | 13747.9372 | 5
count_5 | 3566.3496 | 3548.7479 | 3605.0391 | 3518.5205 | 3680.9667 | 3583.9247 | 154548.6622 | 1.563% | 11% | 4.533% | 11984.3918 | 5
count_6 | 3102.3704 | 3091.5916 | 3126.1745 | 3046.8468 | 3206.2843 | 3114.6535 | 134312.4002 | 1.687% | 11% | 5.119% | 10415.1819 | 5
count_7 | 2692.9628 | 2685.4213 | 2694.0313 | 2648.6643 | 2779.2995 | 2700.0758 | 116434.6741 | 1.591% | 11% | 4.838% | 9028.8634 | 5
count_8 | 2317.602 | 2304.6868 | 2314.262 | 2272.6095 | 2402.5348 | 2322.339 | 100145.6267 | 1.857% | 11% | 5.595% | 7765.7381 | 5
count_9 | 1977.737 | 1973.5635 | 1988.4819 | 1947.1157 | 2045.0324 | 1986.3861 | 85658.4153 | 1.627% | 11% | 4.929% | 6642.3351 | 5
count_10 | 1676.2652 | 1677.6149 | 1685.3125 | 1653.6095 | 1737.6355 | 1686.0875 | 72708.717 | 1.652% | 11% | 4.983% | 5638.1578 | 5
count_11 | 1410.2891 | 1415.5807 | 1409.0223 | 1387.6194 | 1458.3297 | 1416.1682 | 61069.0576 | 1.634% | 11% | 4.993% | 4735.5668 | 5
count_12 | 1168.339 | 1165.3477 | 1164.5718 | 1151.4387 | 1208.3761 | 1171.6147 | 50523.237 | 1.645% | 11% | 4.86% | 3917.7969 | 5
count_13 | 952.4673 | 953.9559 | 956.9216 | 948.6213 | 982.5007 | 958.8934 | 41350.1106 | 1.262% | 11% | 3.533% | 3206.4718 | 5
count_14 | 767.9864 | 766.3957 | 774.3047 | 767.7022 | 788.0415 | 772.8861 | 33328.968 | 1.043% | 11% | 2.801% | 2584.4767 | 5
count_15 | 618.7597 | 611.7188 | 616.0291 | 613.2675 | 615.2525 | 615.0055 | 26520.7237 | 0.3918% | 11% | 1.145% | 2056.5351 | 5
count_16 | 484.7454 | 475.4885 | 479.326 | 486.349 | 479.2686 | 481.0355 | 20743.5693 | 0.8253% | 11% | 2.258% | 1608.5488 | 5
count_17 | 373.9119 | 372.4851 | 370.0789 | 377.66 | 369.0827 | 372.6437 | 16069.4186 | 0.8144% | 11% | 2.302% | 1246.0943 | 5
count_18 | 290.0021 | 290.2198 | 288.8646 | 292.8183 | 286.3 | 289.641 | 12490.112 | 0.7296% | 11% | 2.25% | 968.5389 | 5
count_19 | 221.7876 | 229.4948 | 223.8008 | 226.7794 | 220.8305 | 224.5386 | 9682.7212 | 1.429% | 11.3% | 3.859% | 750.8413 | 5
count_20 | 171.4417 | 177.4775 | 176.5027 | 174.2693 | 171.8717 | 174.3126 | 7516.8367 | 1.382% | 11.3% | 3.463% | 582.889 | 5
count_21 | 137.5156 | 140.0113 | 138.8949 | 139.0715 | 135.6399 | 138.2266 | 5960.7119 | 1.099% | 11.3% | 3.162% | 462.2202 | 5
count_22 | 106.487 | 110.1071 | 108.2089 | 107.5425 | 103.7651 | 107.2221 | 4623.7117 | 1.951% | 11.9% | 5.915% | 358.5432 | 5
count_23 | 81.4952 | 86.5045 | 83.1755 | 83.1219 | 78.7698 | 82.6134 | 3562.5145 | 3.05% | 11.9% | 9.363% | 276.2532 | 5
count_24 | 65.9205 | 68.287 | 66.6788 | 67.0708 | 60.7685 | 65.7451 | 2835.1101 | 3.96% | 11.9% | 11.44% | 219.8471 | 5
count_25 | 48.7764 | 52.0173 | 49.7207 | 51.2489 | 44.9458 | 49.3418 | 2127.7543 | 5.012% | 11.9% | 14.33% | 164.9955 | 5
count_26 | 36.8237 | 40.3306 | 38.4153 | 38.6373 | 32.3334 | 37.3081 | 1608.8268 | 7.302% | 13.4% | 21.44% | 124.7556 | 5
count_27 | 28.6139 | 32.5395 | 29.9939 | 29.9239 | 23.2755 | 28.8693 | 1244.9242 | 10.64% | 13.4% | 32.09% | 96.537 | 5
count_28 | 23.0601 | 23.7171 | 22.8415 | 22.7009 | 17.3133 | 21.9266 | 945.5346 | 10.64% | 15.6% | 29.21% | 73.321 | 5
count_29 | 15.8161 | 18.1029 | 17.0735 | 16.739 | 13.3003 | 16.2064 | 698.8627 | 10.04% | 15.6% | 29.63% | 54.1929 | 5
count_30 | 10.6246 | 13.7491 | 13.0358 | 11.6944 | 9.1726 | 11.6553 | 502.6079 | 14.11% | 15.6% | 39.27% | 38.9745 | 5
count_38 | 7.9684 | 10.4264 | 9.8057 | 8.5988 | 6.1915 | 8.5982 | 370.7764 | 17.24% | 19.3% | 49.25% | 28.7517 | 5
count_50 | 5.7952 | 7.6766 | 7.3831 | 5.8472 | 3.8984 | 6.1201 | 263.915 | 22.08% | 19.3% | 61.73% | 20.4651 | 5
count_70 | 4.2257 | 5.385 | 4.7298 | 4.3567 | 3.0958 | 4.3586 | 187.955 | 17.18% | 27.5% | 52.52% | 14.5749 | 5
Table C.2
| Mean counted X |  | CV_MAX
X >= | X < | DQMAX / SAMPLE | N=3 | N=6 | N=10 | N=20 | N=40
10000 |  | 11 | 7.8 | 5.3 | 4 | 2.7 | 1.8
5000 | 10000 | 11.3 | 7.8 | 5.3 | 4 | 2.7 | 1.9
2000 | 5000 | 11.9 | 7.9 | 5.4 | 4.1 | 2.9 | 2.1
1000 | 2000 | 13.4 | 8 | 5.6 | 4.3 | 3.1 | 2.4
500 | 1000 | 15.6 | 8.2 | 5.9 | 4.7 | 3.6 | 2.9
200 | 500 | 19.3 | 8.9 | 6.8 | 5.7 | 4.7 | 4.1
100 | 200 | 27.5 | 10 | 8 | 7 | 6.1 | 5.5
0 | Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5
NAME\mL | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216
4 | 33813 | 35435 | 35865 | 35225 | 36964 | 33813 | 35435 | 35865 | 35225 | 36964 | 33813 | 35435 | 35865 | 35225 | 36964
5 | 29539 | 30973 | 31250 | 30689 | 32104 | 29539 | 30973 | 31250 | 30689 | 32104 | 29539 | 30973 | 31250 | 30689 | 32104
6 | 25696 | 26983 | 27099 | 26575 | 27964 | 25696 | 26983 | 27099 | 26575 | 27964 | 25696 | 26983 | 27099 | 26575 | 27964
7 | 22305 | 23438 | 23353 | 23102 | 24240 | 22305 | 23438 | 23353 | 23102 | 24240 | 22305 | 23438 | 23353 | 23102 | 24240
8 | 19196 | 20115 | 20061 | 19822 | 20954 | 19196 | 20115 | 20061 | 19822 | 20954 | 19196 | 20115 | 20061 | 19822 | 20954
9 | 16381 | 17225 | 17237 | 16983 | 17836 | 16381 | 17225 | 17237 | 16983 | 17836 | 16381 | 17225 | 17237 | 16983 | 17836
10 | 13884 | 14642 | 14609 | 14423 | 15155 | 13884 | 14642 | 14609 | 14423 | 15155 | 13884 | 14642 | 14609 | 14423 | 15155
11 | 11681 | 12355 | 12214 | 12103 | 12719 | 11681 | 12355 | 12214 | 12103 | 12719 | 11681 | 12355 | 12214 | 12103 | 12719
12 | 9677 | 10171 | 10095 | 10043 | 10539 | 9677 | 10171 | 10095 | 10043 | 10539 | 9677 | 10171 | 10095 | 10043 | 10539
13 | 7889 | 8326 | 8295 | 8274 | 8569 | 7889 | 8326 | 8295 | 8274 | 8569 | 7889 | 8326 | 8295 | 8274 | 8569
14 | 6361 | 6689 | 6712 | 6696 | 6873 | 6361 | 6689 | 6712 | 6696 | 6873 | 6361 | 6689 | 6712 | 6696 | 6873
15 | 5125 | 5339 | 5340 | 5349 | 5366 | 5125 | 5339 | 5340 | 5349 | 5366 | 5125 | 5339 | 5340 | 5349 | 5366
16 | 4015 | 4150 | 4155 | 4242 | 4180 | 4015 | 4150 | 4155 | 4242 | 4180 | 4015 | 4150 | 4155 | 4242 | 4180
17 | 3097 | 3251 | 3208 | 3294 | 3219 | 3097 | 3251 | 3208 | 3294 | 3219 | 3097 | 3251 | 3208 | 3294 | 3219
18 | 2402 | 2533 | 2504 | 2554 | 2497 | 2402 | 2533 | 2504 | 2554 | 2497 | 2402 | 2533 | 2504 | 2554 | 2497
19 | 1837 | 2003 | 1940 | 1978 | 1926 | 1837 | 2003 | 1940 | 1978 | 1926 | 1837 | 2003 | 1940 | 1978 | 1926
20 | 1420 | 1549 | 1530 | 1520 | 1499 | 1420 | 1549 | 1530 | 1520 | 1499 | 1420 | 1549 | 1530 | 1520 | 1499
21 | 1139 | 1222 | 1204 | 1213 | 1183 | 1139 | 1222 | 1204 | 1213 | 1183 | 1139 | 1222 | 1204 | 1213 | 1183
22 | 882 | 961 | 938 | 938 | 905 | 882 | 961 | 938 | 938 | 905 | 882 | 961 | 938 | 938 | 905
23 | 675 | 755 | 721 | 725 | 687 | 675 | 755 | 721 | 725 | 687 | 675 | 755 | 721 | 725 | 687
24 | 546 | 596 | 578 | 585 | 530 | 546 | 596 | 578 | 585 | 530 | 546 | 596 | 578 | 585 | 530
25 | 404 | 454 | 431 | 447 | 392 | 404 | 454 | 431 | 447 | 392 | 404 | 454 | 431 | 447 | 392
26 | 305 | 352 | 333 | 337 | 282 | 305 | 352 | 333 | 337 | 282 | 305 | 352 | 333 | 337 | 282
27 | 237 | 284 | 260 | 261 | 203 | 237 | 284 | 260 | 261 | 203 | 237 | 284 | 260 | 261 | 203
28 | 191 | 207 | 198 | 198 | 151 | 191 | 207 | 198 | 198 | 151 | 191 | 207 | 198 | 198 | 151
29 | 131 | 158 | 148 | 146 | 116 | 131 | 158 | 148 | 146 | 116 | 131 | 158 | 148 | 146 | 116
30 | 88 | 120 | 113 | 102 | 80 | 88 | 120 | 113 | 102 | 80 | 88 | 120 | 113 | 102 | 80
38 | 66 | 91 | 85 | 75 | 54 | 66 | 91 | 85 | 75 | 54 | 66 | 91 | 85 | 75 | 54
50 | 48 | 67 | 64 | 51 | 34 | 48 | 67 | 64 | 51 | 34 | 48 | 67 | 64 | 51 | 34
70 | 35 | 47 | 41 | 38 | 27 | 35 | 47 | 41 | 38 | 27 | 35 | 47 | 41 | 38 | 27

## 6 - B3
Data Preprocessing and Quality Assessment per ISO 11171 Clause 6.6
BOTTLE 3
Serial # | 11171-00009
Date | 2025-11-28
RM | 34D
| Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | X / ml | NTOTAL | CV | DQMAX | DQ | X̅N | NC
Scanned ml | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 43.1227
count_4 | 4082.3649 | 4059.984 | 4137.4312 | 4038.5768 | 4238.2025 | 4111.3119 | 177291.0417 | 1.739% | 11% | 4.856% | 13712.231 | 5
count_5 | 3566.3496 | 3548.7479 | 3605.0391 | 3518.5205 | 3680.9667 | 3583.9247 | 154548.6622 | 1.563% | 11% | 4.533% | 11953.2658 | 5
count_6 | 3102.3704 | 3091.5916 | 3126.1745 | 3046.8468 | 3206.2843 | 3114.6535 | 134312.4002 | 1.687% | 11% | 5.119% | 10388.1315 | 5
count_7 | 2692.9628 | 2685.4213 | 2694.0313 | 2648.6643 | 2779.2995 | 2700.0758 | 116434.6741 | 1.591% | 11% | 4.838% | 9005.4135 | 5
count_8 | 2317.602 | 2304.6868 | 2314.262 | 2272.6095 | 2402.5348 | 2322.339 | 100145.6267 | 1.857% | 11% | 5.595% | 7745.5688 | 5
count_9 | 1977.737 | 1973.5635 | 1988.4819 | 1947.1157 | 2045.0324 | 1986.3861 | 85658.4153 | 1.627% | 11% | 4.929% | 6625.0836 | 5
count_10 | 1676.2652 | 1677.6149 | 1685.3125 | 1653.6095 | 1737.6355 | 1686.0875 | 72708.717 | 1.652% | 11% | 4.983% | 5623.5144 | 5
count_11 | 1410.2891 | 1415.5807 | 1409.0223 | 1387.6194 | 1458.3297 | 1416.1682 | 61069.0576 | 1.634% | 11% | 4.993% | 4723.2675 | 5
count_12 | 1168.339 | 1165.3477 | 1164.5718 | 1151.4387 | 1208.3761 | 1171.6147 | 50523.237 | 1.645% | 11% | 4.86% | 3907.6215 | 5
count_13 | 952.4673 | 953.9559 | 956.9216 | 948.6213 | 982.5007 | 958.8934 | 41350.1106 | 1.262% | 11% | 3.533% | 3198.1439 | 5
count_14 | 767.9864 | 766.3957 | 774.3047 | 767.7022 | 788.0415 | 772.8861 | 33328.968 | 1.043% | 11% | 2.801% | 2577.7642 | 5
count_15 | 618.7597 | 611.7188 | 616.0291 | 613.2675 | 615.2525 | 615.0055 | 26520.7237 | 0.3918% | 11% | 1.145% | 2051.1938 | 5
count_16 | 484.7454 | 475.4885 | 479.326 | 486.349 | 479.2686 | 481.0355 | 20743.5693 | 0.8253% | 11% | 2.258% | 1604.371 | 5
count_17 | 373.9119 | 372.4851 | 370.0789 | 377.66 | 369.0827 | 372.6437 | 16069.4186 | 0.8144% | 11% | 2.302% | 1242.8579 | 5
count_18 | 290.0021 | 290.2198 | 288.8646 | 292.8183 | 286.3 | 289.641 | 12490.112 | 0.7296% | 11% | 2.25% | 966.0234 | 5
count_19 | 221.7876 | 229.4948 | 223.8008 | 226.7794 | 220.8305 | 224.5386 | 9682.7212 | 1.429% | 11.3% | 3.859% | 748.8912 | 5
count_20 | 171.4417 | 177.4775 | 176.5027 | 174.2693 | 171.8717 | 174.3126 | 7516.8367 | 1.382% | 11.3% | 3.463% | 581.3751 | 5
count_21 | 137.5156 | 140.0113 | 138.8949 | 139.0715 | 135.6399 | 138.2266 | 5960.7119 | 1.099% | 11.3% | 3.162% | 461.0197 | 5
count_22 | 106.487 | 110.1071 | 108.2089 | 107.5425 | 103.7651 | 107.2221 | 4623.7117 | 1.951% | 11.9% | 5.915% | 357.612 | 5
count_23 | 81.4952 | 86.5045 | 83.1755 | 83.1219 | 78.7698 | 82.6134 | 3562.5145 | 3.05% | 11.9% | 9.363% | 275.5358 | 5
count_24 | 65.9205 | 68.287 | 66.6788 | 67.0708 | 60.7685 | 65.7451 | 2835.1101 | 3.96% | 11.9% | 11.44% | 219.2761 | 5
count_25 | 48.7764 | 52.0173 | 49.7207 | 51.2489 | 44.9458 | 49.3418 | 2127.7543 | 5.012% | 11.9% | 14.33% | 164.567 | 5
count_26 | 36.8237 | 40.3306 | 38.4153 | 38.6373 | 32.3334 | 37.3081 | 1608.8268 | 7.302% | 13.4% | 21.44% | 124.4316 | 5
count_27 | 28.6139 | 32.5395 | 29.9939 | 29.9239 | 23.2755 | 28.8693 | 1244.9242 | 10.64% | 13.4% | 32.09% | 96.2862 | 5
count_28 | 23.0601 | 23.7171 | 22.8415 | 22.7009 | 17.3133 | 21.9266 | 945.5346 | 10.64% | 15.6% | 29.21% | 73.1305 | 5
count_29 | 15.8161 | 18.1029 | 17.0735 | 16.739 | 13.3003 | 16.2064 | 698.8627 | 10.04% | 15.6% | 29.63% | 54.0522 | 5
count_30 | 10.6246 | 13.7491 | 13.0358 | 11.6944 | 9.1726 | 11.6553 | 502.6079 | 14.11% | 15.6% | 39.27% | 38.8732 | 5
count_38 | 7.9684 | 10.4264 | 9.8057 | 8.5988 | 6.1915 | 8.5982 | 370.7764 | 17.24% | 19.3% | 49.25% | 28.677 | 5
count_50 | 5.7952 | 7.6766 | 7.3831 | 5.8472 | 3.8984 | 6.1201 | 263.915 | 22.08% | 19.3% | 61.73% | 20.412 | 5
count_70 | 4.2257 | 5.385 | 4.7298 | 4.3567 | 3.0958 | 4.3586 | 187.955 | 17.18% | 27.5% | 52.52% | 14.537 | 5
Table C.2
| Mean counted X |  | CV_MAX
X >= | X < | DQMAX / SAMPLE | N=3 | N=6 | N=10 | N=20 | N=40
10000 |  | 11 | 7.8 | 5.3 | 4 | 2.7 | 1.8
5000 | 10000 | 11.3 | 7.8 | 5.3 | 4 | 2.7 | 1.9
2000 | 5000 | 11.9 | 7.9 | 5.4 | 4.1 | 2.9 | 2.1
1000 | 2000 | 13.4 | 8 | 5.6 | 4.3 | 3.1 | 2.4
500 | 1000 | 15.6 | 8.2 | 5.9 | 4.7 | 3.6 | 2.9
200 | 500 | 19.3 | 8.9 | 6.8 | 5.7 | 4.7 | 4.1
100 | 200 | 27.5 | 10 | 8 | 7 | 6.1 | 5.5
0 | Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5 | Portion 1 | Portion 2 | Portion 3 | Portion 4 | Portion 5
NAME\mL | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216 | 8.2827 | 8.7279 | 8.6684 | 8.7221 | 8.7216
4 | 33813 | 35435 | 35865 | 35225 | 36964 | 33813 | 35435 | 35865 | 35225 | 36964 | 33813 | 35435 | 35865 | 35225 | 36964
5 | 29539 | 30973 | 31250 | 30689 | 32104 | 29539 | 30973 | 31250 | 30689 | 32104 | 29539 | 30973 | 31250 | 30689 | 32104
6 | 25696 | 26983 | 27099 | 26575 | 27964 | 25696 | 26983 | 27099 | 26575 | 27964 | 25696 | 26983 | 27099 | 26575 | 27964
7 | 22305 | 23438 | 23353 | 23102 | 24240 | 22305 | 23438 | 23353 | 23102 | 24240 | 22305 | 23438 | 23353 | 23102 | 24240
8 | 19196 | 20115 | 20061 | 19822 | 20954 | 19196 | 20115 | 20061 | 19822 | 20954 | 19196 | 20115 | 20061 | 19822 | 20954
9 | 16381 | 17225 | 17237 | 16983 | 17836 | 16381 | 17225 | 17237 | 16983 | 17836 | 16381 | 17225 | 17237 | 16983 | 17836
10 | 13884 | 14642 | 14609 | 14423 | 15155 | 13884 | 14642 | 14609 | 14423 | 15155 | 13884 | 14642 | 14609 | 14423 | 15155
11 | 11681 | 12355 | 12214 | 12103 | 12719 | 11681 | 12355 | 12214 | 12103 | 12719 | 11681 | 12355 | 12214 | 12103 | 12719
12 | 9677 | 10171 | 10095 | 10043 | 10539 | 9677 | 10171 | 10095 | 10043 | 10539 | 9677 | 10171 | 10095 | 10043 | 10539
13 | 7889 | 8326 | 8295 | 8274 | 8569 | 7889 | 8326 | 8295 | 8274 | 8569 | 7889 | 8326 | 8295 | 8274 | 8569
14 | 6361 | 6689 | 6712 | 6696 | 6873 | 6361 | 6689 | 6712 | 6696 | 6873 | 6361 | 6689 | 6712 | 6696 | 6873
15 | 5125 | 5339 | 5340 | 5349 | 5366 | 5125 | 5339 | 5340 | 5349 | 5366 | 5125 | 5339 | 5340 | 5349 | 5366
16 | 4015 | 4150 | 4155 | 4242 | 4180 | 4015 | 4150 | 4155 | 4242 | 4180 | 4015 | 4150 | 4155 | 4242 | 4180
17 | 3097 | 3251 | 3208 | 3294 | 3219 | 3097 | 3251 | 3208 | 3294 | 3219 | 3097 | 3251 | 3208 | 3294 | 3219
18 | 2402 | 2533 | 2504 | 2554 | 2497 | 2402 | 2533 | 2504 | 2554 | 2497 | 2402 | 2533 | 2504 | 2554 | 2497
19 | 1837 | 2003 | 1940 | 1978 | 1926 | 1837 | 2003 | 1940 | 1978 | 1926 | 1837 | 2003 | 1940 | 1978 | 1926
20 | 1420 | 1549 | 1530 | 1520 | 1499 | 1420 | 1549 | 1530 | 1520 | 1499 | 1420 | 1549 | 1530 | 1520 | 1499
21 | 1139 | 1222 | 1204 | 1213 | 1183 | 1139 | 1222 | 1204 | 1213 | 1183 | 1139 | 1222 | 1204 | 1213 | 1183
22 | 882 | 961 | 938 | 938 | 905 | 882 | 961 | 938 | 938 | 905 | 882 | 961 | 938 | 938 | 905
23 | 675 | 755 | 721 | 725 | 687 | 675 | 755 | 721 | 725 | 687 | 675 | 755 | 721 | 725 | 687
24 | 546 | 596 | 578 | 585 | 530 | 546 | 596 | 578 | 585 | 530 | 546 | 596 | 578 | 585 | 530
25 | 404 | 454 | 431 | 447 | 392 | 404 | 454 | 431 | 447 | 392 | 404 | 454 | 431 | 447 | 392
26 | 305 | 352 | 333 | 337 | 282 | 305 | 352 | 333 | 337 | 282 | 305 | 352 | 333 | 337 | 282
27 | 237 | 284 | 260 | 261 | 203 | 237 | 284 | 260 | 261 | 203 | 237 | 284 | 260 | 261 | 203
28 | 191 | 207 | 198 | 198 | 151 | 191 | 207 | 198 | 198 | 151 | 191 | 207 | 198 | 198 | 151
29 | 131 | 158 | 148 | 146 | 116 | 131 | 158 | 148 | 146 | 116 | 131 | 158 | 148 | 146 | 116
30 | 88 | 120 | 113 | 102 | 80 | 88 | 120 | 113 | 102 | 80 | 88 | 120 | 113 | 102 | 80
38 | 66 | 91 | 85 | 75 | 54 | 66 | 91 | 85 | 75 | 54 | 66 | 91 | 85 | 75 | 54
50 | 48 | 67 | 64 | 51 | 34 | 48 | 67 | 64 | 51 | 34 | 48 | 67 | 64 | 51 | 34
70 | 35 | 47 | 41 | 38 | 27 | 35 | 47 | 41 | 38 | 27 | 35 | 47 | 41 | 38 | 27

## 6 - Thresholds
| X̅N for First Sample | X̅N for Second Sample | X̅N for Third Sample | Mean X̅N | Threshold Voltage Setting | NC | Uncertainty S | Particle Size | Undiluted Particle Concentration | Interpolated Threshold Voltage Setting
| particles per millilitre | particles per millilitre | particles per millilitre | particles per millilitre |  | samples | /ml | µm(c) | /ml
30 | 14.5833 | 14.5749 | 14.537 | 14.5651 | 14000 | 15 | 0.19335 | 4 | 10750.3 | #NAME?
29 | 20.477 | 20.4651 | 20.412 | 20.4514 | 12012.2142 | 15 | 0.34898 | 5 | 7121.1 | #NAME?
28 | 28.7682 | 28.7517 | 28.677 | 28.7323 | 10306.6635 | 15 | 0.38278 | 6 | 4729.1 | #NAME?
27 | 38.9969 | 38.9745 | 38.8732 | 38.9482 | 8843.275 | 15 | 0.42454 | 7 | 3199.3 | #NAME?
26 | 54.2242 | 54.1929 | 54.0522 | 54.1564 | 7587.6652 | 15 | 0.42 | 8 | 2280.6 | #NAME?
25 | 73.3633 | 73.321 | 73.1305 | 73.2716 | 6510.3328 | 15 | 0.60233 | 9 | 1615 | #NAME?
24 | 96.5927 | 96.537 | 96.2862 | 96.472 | 5585.9652 | 15 | 0.79333 | 10 | 1130.5 | #NAME?
23 | 124.8276 | 124.7556 | 124.4316 | 124.6716 | 4792.8436 | 15 | 0.70342 | 11 | 816.2 | #NAME?
22 | 165.0907 | 164.9955 | 164.567 | 164.8844 | 4112.3331 | 15 | 0.63858 | 12 | 618.6 | #NAME?
21 | 219.9739 | 219.8471 | 219.2761 | 219.699 | 3528.4447 | 15 | 0.67219 | 13 | 481.9 | #NAME?
20 | 276.4126 | 276.2532 | 275.5358 | 276.0672 | 3027.4595 | 15 | 0.65053 | 14 | 384.7 | #NAME?
19 | 358.7501 | 358.5432 | 357.612 | 358.3018 | 2597.6066 | 15 | 0.54025 | 15 | 314.6 | #NAME?
18 | 462.4868 | 462.2202 | 461.0197 | 461.9089 | 2228.7862 | 15 | 0.3923 | 16 | 261.8 | #NAME?
17 | 583.2253 | 582.889 | 581.3751 | 582.4965 | 1912.3327 | 15 | 0.62199 | 17 | 221.2 | #NAME?
16 | 751.2746 | 750.8413 | 748.8912 | 750.3357 | 1640.8107 | 15 | 0.82823 | 18 | 188.7 | #NAME?
15 | 969.0978 | 968.5389 | 966.0234 | 967.8867 | 1407.8407 | 15 | 0.5456 | 19 | 161 | #NAME?
14 | 1246.8133 | 1246.0943 | 1242.8579 | 1245.2552 | 1207.9488 | 15 | 0.78363 | 20 | 136.4 | #NAME?
13 | 1609.4769 | 1608.5488 | 1604.371 | 1607.4656 | 1036.4386 | 15 | 1.0251 | 21 | 114.4 | #NAME?
12 | 2057.7216 | 2056.5351 | 2051.1938 | 2055.1502 | 889.2802 | 15 | 0.6221 | 22 | 95.3 | #NAME?
11 | 2585.9678 | 2584.4767 | 2577.7642 | 2582.7363 | 763.016 | 15 | 2.0807 | 23 | 79 | #NAME?
10 | 3208.3219 | 3206.4718 | 3198.1439 | 3204.3125 | 654.6794 | 15 | 3.1249 | 24 | 65.3 | #NAME?
9 | 3920.0574 | 3917.7969 | 3907.6215 | 3915.1586 | 561.7249 | 15 | 4.9774 | 25 | 53.8 | #NAME?
8 | 4738.2991 | 4735.5668 | 4723.2675 | 4732.3778 | 481.9686 | 15 | 5.9762 | 26 | 44.6 | #NAME?
7 | 5641.4109 | 5638.1578 | 5623.5144 | 5634.361 | 413.5364 | 15 | 7.1934 | 27 | 37 | #NAME?
6 | 6646.1676 | 6642.3351 | 6625.0836 | 6637.8621 | 354.8206 | 15 | 8.3468 | 28 | 30.7 | #NAME?
5 | 7770.2187 | 7765.7381 | 7745.5688 | 7760.5085 | 304.4415 | 15 | 11.1377 | 29 | 25.2 | #NAME?
4 | 9034.0728 | 9028.8634 | 9005.4135 | 9022.7832 | 261.2154 | 15 | 11.0884 | 30 | 20.1 | #NAME?
3 | 10421.1912 | 10415.1819 | 10388.1315 | 10408.1682 | 224.1268 | 15 | 13.5693
2 | 11991.3065 | 11984.3918 | 11953.2658 | 11976.3214 | 192.3043 | 15 | 14.4623
1 | 13755.8694 | 13747.9372 | 13712.231 | 13738.6792 | 165 | 15 | 18.4558