## Annex A
Annex A Preliminary Check
Analysis using Bottle Samples and 25% coincidence RM8632a | [8 empty] | Concentrate Dilution Calculator |  |  | Concentrate Manufacturing Calculator
Can perform Annex C at same time using 5x bottles |  |  |  | Volumetric Sampling Accuracy |  |  | Volume Actually Measured |  | Concentrate | 100 | mg/l | Target Final Fluid Volume ml | 250
[9 empty] | Final fluid volume | 100 | ml | Target Concentration mg/l | 1000
APC Model |  |  | LPA3-NAVAIR |  |  |  | Prepare 5 bottles of 10um PSL spheres of concentration 2000/ml |  | Target concentration | 1 | mg/l | mg required | 250
APC S/N |  |  | 11171-00009 |  |  |  | Obtain counts using 10ml flush + 5x10ml portions
Sensor Model |  |  | LPA3-NAVAIR |  |  |  | Import all portion data into following sheet (A - B1) |  | Concentrate to add | 1 | ml
Sensor S/N |  |  | 11171-00009 |  |  |  | Adjust A - B1 to A-B5 sheets deleting outliers as per 6.6 |  | Diluent to add | 99 | ml
Date |  |  | 2026-06-08 |  |  |  | Enter expected concentration in cell below i.e. 2000
Operator |  |  | JMD |  |  |  | Can also export histogram data for one bottle and perform Annex D
Threshold Noise |  |  | 35 | [5 empty] | Alt Concentrate Dilution Calculator
Photocurrent |  |  | 12236 | [5 empty] | Concentrate mg/l | 100
4um(c) Coincidence Limit /ml |  |  | 20000 | [5 empty] | Diluent volume ml | 500
A.1 RM8632 4um count/mg/l |  |  | 6300 | [5 empty] | Target concentration mg/l | 6.5
RM8632x mg/l at 25% coincidence limit |  |  | 0.79365 | OR use measuring cylinder to weigh 100ml into a bottle |  |  |  |  | Concentrate to add | 32.5
[9 empty] | Added | 3.27
Concentrate mg/l |  |  | 100 | [5 empty] | Final Concentration | 0.64975
Wanted fluid volume ml |  |  | 100
Target quantity of dust in dilution |  |  | 0.079365
Concentrate to add ml |  |  | 0.79365
Clean fluid to add |  |  | 99.2063 | Date | 2026-06-08 |  |  |  | Threshold factor | 4.4583
|  |  |  | Operator | JMD |  |  |  |  | 160
|  | /ml | Deviation | Particle Counter | 11171-00009 |  | Date |  |  | 713.3357
| Portion 1 | 6385.56 | -1.714% |  |  |  | Operator | JMD
| Portion 2 | 6535.61 | 0.5953% | Measure Fluid Density |  |  | Particle Counter | 11171-00001
| Portion 3 | 6581.8 | 1.306% | Temperature | 40'C
| Portion 4 | 6506.8 | 0.1519% | 100ml Weight | 85.12 |  | AVERAGE 10um count/ml | 2598
| Portion 5 | 6474.9 | -0.3391% | Density g/cc | 0.8512 |  | Expected 10um count/ml | 2000
[7 empty] | Capture Fraction | 1.299
| Mean | 6496.934 |  | Full bottle weight | 178.15
| Sample StdDev | 65.856 |  | Flush | 20 | ml | Set "Beam Flow Correction" | 299
|  |  |  | Sample | 50 | ml
|  |  |  | Total Fluid Programmed | 70 | ml
|  |  |  | Final Measured Weight | 117.85 | g
|  |  |  | Fluid Weight Taken | 60.3
|  |  |  | ml taken | 70.8412
| CVVOL | 1.014% | <3% = PASS | Error | 1.202%
Used microbalance to weigh 10ml sample of fluid, added using precision digital pipette. Added 11x 10ml I total to plastic sample bottle and weighed result (after taring empty bottle weight).
Used bottle sampler to sample 20ml pumped flush + 50ml pumped sample.
Measured final weight of bottle+remaining fluid.