## BS110 Cal'd Thresholds
Cal'd thresholds | [10 empty] | Counts
S/N | Ref material | 4 | 6 | 10 | 14 | 21 | 25 | 38 | 50 | 70 | 4 | 6 | 10 | 14 | 21 | 25 | 38 | 50 | 70
63010066 | 37D | 78 | 186 | 522 | 1012 | 2072 | 3097 | 10183 | 15921 | 24781 | 1057464 | 459753 | 104698 | 1012 | 2072 | 4776 | - | - | -
63010067 | 37D | 76 | 180 | 508 | 994 | 2074 | 3050 | 9964 | 16287 | 25144 | 1064210 | 459529 | 106782 | 36466 | 10948 | 5210 | - | - | -
[8 empty] | 10301 | 15776 | 24523
[8 empty] | 9887 | 16401 | 25376
[8 empty] | 10126 | 15693 | 24912
[8 empty] | 10249 | 16114 | 25502
[8 empty] | 9926 | 15882 | 24367
[8 empty] | 10314 | 16352 | 25228
[8 empty] | 10041 | 15731 | 24689
[8 empty] | 9853 | 16228 | 25551
[8 empty] | 10217 | 16276 | 24441
[8 empty] | 9874 | 15648 | 25061
[8 empty] | 10296 | 16082 | 25407
[8 empty] | 10008 | 16418 | 24602
[8 empty] | 10154 | 15872 | 25309
[8 empty] | 9938 | 16215 | 24388
[8 empty] | 10311 | 15734 | 25536
[8 empty] | 9840 | 16332 | 24723
[8 empty] | 10187 | 15905 | 25124
[8 empty] | 10058 | 16391 | 24467
[8 empty] | 10072 | 16030 | 24934
Step 1 |  | Prepare PartiStan - 60s shake, 60s ultrasonic, 60s turn over/roll bottle, 60s ultrasonic
Step 2 |  | Decant into 4x bottles (Cleaned with heptane + dried)
Step 3 |  | Open LPA View, connect LPA3 via USB-NETTWarrior to laptop. Ctrl+N > LPA3 Calibration > Select LPA3 COM port
Step 4 |  | Select PartiStan varient (EG, Partistan 2806 Lot 35D.) Click APPLY Edit
Step 5 |  | Shake 1x decanted bottle of fluid for 60s. Remove lid & put bottle into bottle sampler. Fit Vacuum cap 60s Vacuum. If bubbles still appear at top, another 15-30s vacuum.
Step 6 |  | switch bottle sampler cap - set to unit to sample. Turn on unit. Let bottle sampler unit it reaches pressure and stops
Step 7 |  | Start test on LPA3 once Bottle sampler switches off.
Step 8 |  | Allow LPA3 test to run
Step 9 |  | Once LPA3 says EMPTYING, click 'AUTO ADJUST' button on LPA view. allow 30~s for adjustment to complete.
Step 9 |  | Manually set 38, 50, 70 micron levels on the LPA3 itself. Then on the same screen, click 'recalculate'
Step 10 |  | Once test is finished, repeat steps 4-8. Await test results to auto print to compare as a verification test.

## Sheet1
|  | 4 | 6 | 10 | 14 | 21 | 25 | 38 | 50 | 70
Serial Number 63010045 | 63010045 | -0.821% | -0.1702% | -0.6935% | -0.7332% | -0.05464% | 0.3647%
Serial Number 63010046 | 63010046 | 1.37% | -0.1842% | 0.4868% | -1.888% | -4.153% | -5.701%
Serial Number 63010047 | 63010047 | -3.113% | -3.602% | -4.873% | -3.967% | 1.393% | 3.596%
Serial Number 63010048 | 63010048 | 0.3047% | 0.1181% | 0.7689% | 2.074% | 2.851% | 0.07678%
Serial Number 63010049 | 63010049 | -1.226% | -0.7752% | -1.661% | -5.611% | -7.814% | -13.68%
Serial Number 63010050 | 63010050 | -0.3744% | 2.267% | 5.976% | 10.2% | 24.97% | 34.08%
Serial Number 63010051 | 63010051 | -0.047% | -0.07681% | 1.127% | 2.2% | -7.741% | 4.395%
Serial Number 63010052 | 63010052 | -0.09635% | 1.653% | 2.641% | 0.6082% | -5.493% | 0.7456%
Serial Number 63010053 | 63010053 | -3.051% | -2.598% | -3.838% | -3.436% | -7.787% | -8.157%
Serial Number 63010054 | 63010054 | -0.9424% | -0.367% | -0.228% | 0.5096% | -0.09671% | -3.728%
Serial Number 63010055 | 63010055 | 1.584% | 0.732% | 0.01862% | 0.8263% | 1.758% | 4.856%
Serial Number 63010056 | 63010056 | 1.021% | -0.3678% | 0.2243% | -0.05479% | 0.07737% | -0.1316%
Serial Number 63010057 | 63010057 | 0.2176% | 1.103% | -2.814% | 0.8384% | 6.77% | 6.228%
Serial Number 63010058 | 63010058 | -1.264% | -2.659% | -0.1303% | -4.741% | -5.792% | -11.21%
Serial Number 63010059 | 63010059 | -1.216% | 0.1941% | 1.694% | 2.985% | 9.854% | 13.8%
Serial Number 63010060 | 63010060 | -0.7311% | -0.2937% | -0.2709% | 0.1094% | 0.03643% | 0.09597%
Serial Number 63010061 | 63010061 | -2.442% | -3.593% | -0.228% | 4.745% | 5.957% | -3.246%
Serial Number 63010062 | 63010062 | -5.028% | -5.53% | -6.581% | -5.71% | -5.282% | -2.668%
Serial Number 63010063 | 63010063 | 0.04061% | 0.5762% | 0.09308% | -0.05198% | 0.07286% | 0%
Serial Number 63010064 | 63010064 | 0.8895% | 2.78% | 1.214% | -0.1644% | 7.389% | 6.009%
Serial Number 63010065 | 63010065
Serial Number 63010066 | 63010066
Serial Number 63010067 | 63010067
Serial Number 63010068 | 63010068
Serial Number 63010069 | 63010069
Serial Number 63010070 | 63010070
Serial Number 63010071 | 63010071
Serial Number 63010072 | 63010072
Serial Number 63010073 | 63010073
Serial Number 63010074 | 63010074
Serial Number 63010075 | 63010075
Serial Number 63010076 | 63010076
Serial Number 63010077 | 63010077
Serial Number 63010078 | 63010078
Serial Number 63010079 | 63010079
Serial Number 63010080 | 63010080
Serial Number 63010081 | 63010081
Serial Number 63010082 | 63010082

## Serial Number 63010045
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010045 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10720.65 |  | 4 | 10809.4 | 92
6 | 4647.7 | 4639.79 |  | 6 | 4647.7 | 212
10 | 1074.3 | 1066.85 |  | 10 | 1074.3 | 578
14 | 365.5 | 362.82 |  | 14 | 365.5 | 1116
21 | 109.8 | 109.74 |  | 21 | 109.8 | 2204
25 | 52.1 | 52.29 |  | 25 | 52.1 | 3170
38 | - | 1.53 |  | 38 | - | 9921
50 | - | 0.05 |  | 50 | - | 1589
70 | - | 0 |  | 70 | - | 24883
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010046
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010046 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10957.53 |  | 4 | 10809.4 | 90
6 | 4647.7 | 4639.14 |  | 6 | 4647.7 | 218
10 | 1074.3 | 1079.53 |  | 10 | 1074.3 | 600
14 | 365.5 | 358.6 |  | 14 | 365.5 | 1178
21 | 109.8 | 105.24 |  | 21 | 109.8 | 2406
25 | 52.1 | 49.13 |  | 25 | 52.1 | 3528
38 | - | 2.92 |  | 38 | - | 9952
50 | - | 0.71 |  | 50 | - | 16091
70 | - | 0.11 |  | 70 | - | 24765
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010047
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010047 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5641.37 |  | 4 | 5822.6 | 84
6 | 2422.5 | 2335.24 |  | 6 | 2422.5 | 194
10 | 543.8 | 517.3 |  | 10 | 543.8 | 544
14 | 182.5 | 175.26 |  | 14 | 182.5 | 1102
21 | 51.7 | 52.42 |  | 21 | 51.7 | 2350
25 | 22.8 | 23.62 |  | 25 | 22.8 | 3700
38 | - | 2.13 |  | 38 | - | 10245
50 | - | 0.5 |  | 50 | - | 15924
70 | - | 0.11 |  | 70 | - | 24995
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010048
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010048 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10842.34 |  | 4 | 10809.4 | 80
6 | 4647.7 | 4653.19 |  | 6 | 4647.7 | 184
10 | 1074.3 | 1082.56 |  | 10 | 1074.3 | 514
14 | 365.5 | 373.08 |  | 14 | 365.5 | 1006
21 | 109.8 | 112.93 |  | 21 | 109.8 | 2032
25 | 52.1 | 52.14 |  | 25 | 52.1 | 2962
38 | - | 1.4 |  | 38 | - | 10165
50 | - | 0.35 |  | 50 | - | 16205
70 | - | 0.05 |  | 70 | - | 24750
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010049
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010049 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5751.2 |  | 4 | 5822.6 | 80
6 | 2422.5 | 2403.72 |  | 6 | 2422.5 | 190
10 | 543.8 | 534.77 |  | 10 | 543.8 | 542
14 | 182.5 | 172.26 |  | 14 | 182.5 | 1084
21 | 51.7 | 47.66 |  | 21 | 51.7 | 2606
25 | 22.8 | 19.68 |  | 25 | 22.8 | 3570
38 | - | 1.24 |  | 38 | - | 9907
50 | - | 0.34 |  | 50 | - | 16076
70 | - | 0 |  | 70 | - | 24844
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010050
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010050 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5800.8 |  | 4 | 5822.6 | 88
6 | 2422.5 | 2477.42 |  | 6 | 2422.5 | 198
10 | 543.8 | 576.3 |  | 10 | 543.8 | 534
14 | 182.5 | 201.11 |  | 14 | 182.5 | 1022
21 | 51.7 | 64.61 |  | 21 | 51.7 | 2070
25 | 22.8 | 30.57 |  | 25 | 22.8 | 3086
38 | - | 1.59 |  | 38 | - | 10132
50 | - | 0.51 |  | 50 | - | 15873
70 | - | 0 |  | 70 | - | 25054
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010051
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010051 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10804.32 |  | 4 | 10809.4 | 92
6 | 4647.7 | 4644.13 |  | 6 | 4647.7 | 212
10 | 1074.3 | 1086.41 |  | 10 | 1074.3 | 578
14 | 365.5 | 373.54 |  | 14 | 365.5 | 1110
21 | 109.8 | 101.3 |  | 21 | 109.8 | 2210
25 | 52.1 | 54.39 |  | 25 | 52.1 | 3204
38 | - | 5.4 |  | 38 | - | 10011
50 | - | 0.83 |  | 50 | - | 16174
70 | - | 0.23 |  | 70 | - | 24822
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010052
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010052 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5816.99 |  | 4 | 5822.6 | 70
6 | 2422.5 | 2462.54 |  | 6 | 2422.5 | 176
10 | 543.8 | 558.16 |  | 10 | 543.8 | 538
14 | 182.5 | 183.61 |  | 14 | 182.5 | 1074
21 | 51.7 | 48.86 |  | 21 | 51.7 | 2394
25 | 22.8 | 22.97 |  | 25 | 22.8 | 3666
38 | - | 1.78 |  | 38 | - | 10201
50 | - | 0.44 |  | 50 | - | 16053
70 | - | 0.11 |  | 70 | - | 25029
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010053
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010053 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10479.57 |  | 4 | 10809.4 | 88
6 | 4647.7 | 4526.96 |  | 6 | 4647.7 | 198
10 | 1074.3 | 1033.07 |  | 10 | 1074.3 | 546
14 | 365.5 | 352.94 |  | 14 | 365.5 | 1060
21 | 109.8 | 101.25 |  | 21 | 109.8 | 2192
25 | 52.1 | 47.85 |  | 25 | 52.1 | 3290
38 | - | 2.51 |  | 38 | - | 10037
50 | - | 0.54 |  | 50 | - | 16014
70 | - | 0.13 |  | 70 | - | 25093
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010054
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010054 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5767.73 |  | 4 | 5822.6 | 86
6 | 2422.5 | 2413.61 |  | 6 | 2422.5 | 204
10 | 543.8 | 542.56 |  | 10 | 543.8 | 572
14 | 182.5 | 183.43 |  | 14 | 182.5 | 1108
21 | 51.7 | 51.65 |  | 21 | 51.7 | 2382
25 | 22.8 | 21.95 |  | 25 | 22.8 | 3608
38 | - | 1.96 |  | 38 | - | 10029
50 | - | 0.11 |  | 50 | - | 16221
70 | - | 0 |  | 70 | - | 24748
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010055
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010055 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10980.65 |  | 4 | 10809.4 | 82
6 | 4647.7 | 4681.72 |  | 6 | 4647.7 | 190
10 | 1074.3 | 1074.5 |  | 10 | 1074.3 | 530
14 | 365.5 | 368.52 |  | 14 | 365.5 | 1030
21 | 109.8 | 111.73 |  | 21 | 109.8 | 2076
25 | 52.1 | 54.63 |  | 25 | 52.1 | 3036
38 | - | 1.95 |  | 38 | - | 10144
50 | - | 0.37 |  | 50 | - | 15958
70 | - | 0.04 |  | 70 | - | 25096
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010056
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010056 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5882.03 |  | 4 | 5822.6 | 82
6 | 2422.5 | 2413.59 |  | 6 | 2422.5 | 194
10 | 543.8 | 545.02 |  | 10 | 543.8 | 552
14 | 182.5 | 182.4 |  | 14 | 182.5 | 1132
21 | 51.7 | 51.74 |  | 21 | 51.7 | 2460
25 | 22.8 | 22.77 |  | 25 | 22.8 | 3678
38 | - | 2.13 |  | 38 | - | 9988
50 | - | 0.22 |  | 50 | - | 16167
70 | - | 0 |  | 70 | - | 24783
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010057
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010057 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan 28DR
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5835.27 |  | 4 | 5822.6 | 74
6 | 2422.5 | 2449.23 |  | 6 | 2422.5 | 184
10 | 543.8 | 528.5 |  | 10 | 543.8 | 599
14 | 182.5 | 184.03 |  | 14 | 182.5 | 1020
21 | 51.7 | 55.2 |  | 21 | 51.7 | 2080
25 | 22.8 | 24.22 |  | 25 | 22.8 | 2938
38 | - | 1.76 |  | 38 | - | 10084
50 | - | 0.51 |  | 50 | - | 16139
70 | - | 0.11 |  | 70 | - | 25112
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010058
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010058 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10672.8 |  | 4 | 10809.4 | 96
6 | 4647.7 | 4524.14 |  | 6 | 4647.7 | 234
10 | 1074.3 | 1072.9 |  | 10 | 1074.3 | 644
14 | 365.5 | 348.17 |  | 14 | 365.5 | 1252
21 | 109.8 | 103.44 |  | 21 | 109.8 | 2520
25 | 52.1 | 46.26 |  | 25 | 52.1 | 3710
38 | - | 4.38 |  | 38 | - | 10198
50 | - | 0.85 |  | 50 | - | 15876
70 | - | 0.04 |  | 70 | - | 25071
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010059
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010059 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10677.97 |  | 4 | 10809.4 | 90
6 | 4647.7 | 4656.72 |  | 6 | 4647.7 | 204
10 | 1074.3 | 1092.5 |  | 10 | 1074.3 | 552
14 | 365.5 | 376.41 |  | 14 | 365.5 | 1064
21 | 109.8 | 120.62 |  | 21 | 109.8 | 2092
25 | 52.1 | 59.29 |  | 25 | 52.1 | 3036
38 | - | 2.8 |  | 38 | - | 9879
50 | - | 0.59 |  | 50 | - | 16188
70 | - | 0.05 |  | 70 | - | 24977
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010060
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010060 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10730.37 |  | 4 | 10809.4 | 68
6 | 4647.7 | 4634.05 |  | 6 | 4647.7 | 152
10 | 1074.3 | 1071.39 |  | 10 | 1074.3 | 428
14 | 365.5 | 365.9 |  | 14 | 365.5 | 876
21 | 109.8 | 109.84 |  | 21 | 109.8 | 1886
25 | 52.1 | 52.15 |  | 25 | 52.1 | 2856
38 | - | 2.45 |  | 38 | - | 10256
50 | - | 0.28 |  | 50 | - | 15907
70 | - | 0.05 |  | 70 | - | 24810
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010061
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010061 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5680.42 |  | 4 | 5822.6 | 90
6 | 2422.5 | 2335.47 |  | 6 | 2422.5 | 206
10 | 543.8 | 542.56 |  | 10 | 543.8 | 566
14 | 182.5 | 191.16 |  | 14 | 182.5 | 1106
21 | 51.7 | 54.78 |  | 21 | 51.7 | 2376
25 | 22.8 | 22.06 |  | 25 | 22.8 | 3696
38 | - | 1.46 |  | 38 | - | 10177
50 | - | 0.33 |  | 50 | - | 15895
70 | - | 0.05 |  | 70 | - | 25088
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010062
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010062 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10265.89 |  | 4 | 10809.4 | 92
6 | 4647.7 | 4390.66 |  | 6 | 4647.7 | 206
10 | 1074.3 | 1003.6 |  | 10 | 1074.3 | 562
14 | 365.5 | 344.63 |  | 14 | 365.5 | 1092
21 | 109.8 | 104 |  | 21 | 109.8 | 2204
25 | 52.1 | 50.71 |  | 25 | 52.1 | 3178
38 | - | 2.91 |  | 38 | - | 10118
50 | - | 0.52 |  | 50 | - | 15862
70 | - | 0.05 |  | 70 | - | 24791
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010063
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010063 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10813.79 |  | 4 | 10809.4 | 86
6 | 4647.7 | 4674.48 |  | 6 | 4647.7 | 196
10 | 1074.3 | 1075.3 |  | 10 | 1074.3 | 546
14 | 365.5 | 365.31 |  | 14 | 365.5 | 1076
21 | 109.8 | 109.88 |  | 21 | 109.8 | 2230
25 | 52.1 | 52.1 |  | 25 | 52.1 | 3408
38 | - | 3.97 |  | 38 | - | 10233
50 | - | 1.46 |  | 50 | - | 16112
70 | - | 0.4 |  | 70 | - | 25018
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010064
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010064 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan 28DR
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5874.39 |  | 4 | 5822.6 | 84
6 | 2422.5 | 2489.84 |  | 6 | 2422.5 | 198
10 | 543.8 | 550.4 |  | 10 | 543.8 | 562
14 | 182.5 | 182.2 |  | 14 | 182.5 | 1126
21 | 51.7 | 55.52 |  | 21 | 51.7 | 2374
25 | 22.8 | 24.17 |  | 25 | 22.8 | 3670
38 | - | 1.97 |  | 38 | - | 9944
50 | - | 0.67 |  | 50 | - | 15971
70 | - | 0.22 |  | 70 | - | 24856
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.