## CAL (hidden)
CALIBRATION RESULTS
verification_count_0_0 | 3502561
verification_count_0_1 | 1384415
verification_count_0_2 | 97621
verification_count_0_3 | 23939
verification_count_0_4 | 11133
verification_count_0_5 | 1551
verification_count_0_6 | 593
verification_count_0_7 | 71
verification_count_1_0 | 3504210
verification_count_1_1 | 1382038
verification_count_1_2 | 99179
verification_count_1_3 | 24541
verification_count_1_4 | 11031
verification_count_1_5 | 1458
verification_count_1_6 | 482
verification_count_1_7 | 71
verification_count_2_0 | 3524063
verification_count_2_1 | 1391154
verification_count_2_2 | 98798
verification_count_2_3 | 24709
verification_count_2_4 | 11141
verification_count_2_5 | 1411
verification_count_2_6 | 552
verification_count_2_7 | 42
no_noise_threshold | 61
led_level | 11559
CALIBRATION DOCUMENT VALUES
date_ | 2025-12-08 16:48
serial_number_ | 204
RIG PARAMETERS
meas_init_level | 12.57
calc_init_volume | 12.57
desired_concentration | 16
calc_qty_add | 201.12
meas_qty_add | 201.06
meas_final_level | 12.57
calc_volume | 12.57
calc_concentration | 15.9952
rig_param_volume_mult | -1
rig_param_volume_constant | 0
rig_param_identifier | 367
rig_cleanliness_count | 0
ICM SETTINGS
time | 1765210805
serial_number | 204
hwRevision | 0
user_lock | 0
user_pass | 0
CycleCount | 0
last_download | 0
FACTORY SET OPTIONS
SET_OPT_LOG | 1
SET_OPT_WATER | 1
SET_OPT_RELAYS | 1
SET_OPT_K_DISABLED | 0
SET_OPT_CAN | 1
SET_OPT_CAL_DUE_UNSET | 1
SET_OPT_START_SIGNAL_DELIMITS_TEST | 1
SET_OPT_USB_HOST | 1
SET_OPT_4_20mA | 1
SET_OPT_ISO4406_GTE | 0
SET_OPT_ENFORCE_COUNT_FLOOR | 0
SET_OPT_SIMULATED_CALIBRATIONS | 0
FLOW CELL PARAMETERS
effective_slit_width | 68
effective_slit_height | 160
divergence_milliradians | 0
channel_depth | 400
flow_multiplier | 1000
calibration_led_level_initial | 10558
calibration_led_level_last | 109.48
calibration_led_level_alarm | 50
calibration_date_last | 1765210840
calibration_date_next | 1796746840
THRESHOLDS
level[0] | 134
level[1] | 324
level[2] | 1886
level[3] | 4122
level[4] | 6566
level[5] | 16000
level[6] | 22008
level[7] | 32390
source_drive | 4095
detection_threshold | 0
tz_offset_seconds | 0
backlight_rgb[0] | 15
backlight_rgb[1] | 15
backlight_rgb[2] | 15
USER SETTINGS
ref | CALIBRATED
continous_mode_confirm_target_n | 1
ignore_initial_n | 0
result_format | 0
test_time | 120
test_number | 12
test_interval | 0
log_interval | 0
TEST MODE
SET_MODE_CYCLE_CONTINUOUS | 0
SET_MODE_START_TEST_AUTOMATICALLY | 0
SET_MODE_CONTINUOUS_STOP_WHEN_CLEAN | 0
SET_MODE_CONTINUOUS_LOG_EVERY_TEST | 0
SET_MODE_CONTINUOUS_CONFIRM_TARGET | 0
SET_MODE_START_SIGNAL_DELIMITS_TEST | 0
SET_MODE_AS4059E2_NO_4U | 0
SET_MODE_SIMULATE | 0
SET_MODE_LOW_FLOW_CLEAN_DISABLED | 1
SET_MODE_BUBBLE_MITIGATION | 0
SET LIMITS
clean[0] | -2147483648
clean[1] | -2147483648
clean[2] | -2147483648
clean[3] | -2147483648
clean[4] | -2147483648
clean[5] | -2147483648
clean[6] | -2147483648
clean[7] | -2147483648
dirty[0] | -2147483648
dirty[1] | -2147483648
dirty[2] | -2147483648
dirty[3] | -2147483648
dirty[4] | -2147483648
dirty[5] | -2147483648
dirty[6] | -2147483648
dirty[7] | -2147483648
h2o_max | 0
h2o_min | 0
temperature_max | 0
temperature_min | 0
alarm_mode | 0
alarm_mode_2 | 0
modbus_address | 4
can_base_address | 419411263
can_baud | 3
com_options | 0
RH SENSOR CALIBRATION
version | 7
cal_count | 1
serial | 1212607
point[0].rh | 11.666
point[0].capacitance_pf | 139.755
point[0].resistance_ohms | 109.431
point[0].temperature | 24.2238
point[1].rh | 84.2542
point[1].capacitance_pf | 156.852
point[1].resistance_ohms | 109.205
point[1].temperature | 23.6337
cal[0].date | 1659714124
cal[0].calibrator_serial | 51424
cal[0].c0 | 144.032
cal[0].s | 0.23562
cal[0].r0 | 0
cal[0].rs | 0
indicated_rh[0] | 11.4745
indicated_temperature[0] | 25.4883
cal[1].date | 1659714124
cal[1].calibrator_serial | 51424
cal[1].c0 | 144.032
cal[1].s | 0.23562
cal[1].r0 | 0
cal[1].rs | 0
indicated_rh[1] | 83.9265
indicated_temperature[1] | 24.8777
ICM STATUS
firmware_revision | 3.55
clock_battery_voltage_mV | 0

## mtd (hidden)
SRM2806 2.8mg MTD / ml |  |  |  |  | Interpolation Order
[5 empty] | 4 |  | to 10 micron |  | to 70 micron
um | n/ml | n/mg/ml | Expanded Uncertainty | ln(n/mg/ml) | um | n/mg/ml | x | y | x | y
1 | 108400 | 38714.2857 | 0.274 | 10.564 | 50 | #NAME? | 1 | #NAME? | 1 | #NAME?
2 | 27035 | 9655.3571 | 0.265 | 9.1753 |  |  | 1.1 | #NAME? | 2 | #NAME?
3 | 11209 | 4003.2143 | 0.272 | 8.2949 |  |  | 1.2 | #NAME? | 3 | #NAME?
4 | 6095 | 2176.7857 | 0.308 | 7.6856 |  |  | 1.3 | #NAME? | 4 | #NAME?
5 | 3737 | 1334.6429 | 0.414 | 7.1964 | n/mg/l | 10000 | 1.4 | #NAME? | 5 | #NAME?
6 | 2395 | 855.3571 | 0.56 | 6.7515 | ln(n/mg/l) | 9.2103 | 1.5 | #NAME? | 6 | #NAME?
7 | 1573 | 561.7857 | 0.542 | 6.3311 | um | #NAME? | 1.6 | #NAME? | 7 | #NAME?
8 | 1055 | 376.7857 | 0.499 | 5.9317 |  |  | 1.7 | #NAME? | 8 | #NAME?
9 | 725.8 | 259.2143 | 0.544 | 5.5577 |  |  | 1.8 | #NAME? | 9 | #NAME?
10 | 513.7 | 183.4643 | 0.643 | 5.212 |  |  | 1.9 | #NAME? | 10 | #NAME?
11 | 374.6 | 133.7857 | 0.759 | 4.8962 | 2 | #NAME? | 2 | #NAME? | 11 | #NAME?
12 | 281 | 100.3571 | 0.981 | 4.6087 |  |  | 2.1 | #NAME? | 12 | #NAME?
13 | 216.4 | 77.2857 | 1.273 | 4.3475 |  |  | 2.2 | #NAME? | 13 | #NAME?
14 | 170.4 | 60.8571 | 1.655 | 4.1085 | #NAME? |  | 2.3 | #NAME? | 14 | #NAME?
15 | 136.8 | 48.8571 | 2.211 | 3.8889 |  |  | 2.4 | #NAME? | 15 | #NAME?
16 | 111.3 | 39.75 | 2.414 | 3.6826 | (Testing area only) |  | 2.5 | #NAME? | 16 | #NAME?
17 | 91.33 | 32.6179 | 2.91 | 3.4849 |  |  | 2.6 | #NAME? | 17 | #NAME?
18 | 75.29 | 26.8893 | 3.444 | 3.2917 |  |  | 2.7 | #NAME? | 18 | #NAME?
19 | 62.17 | 22.2036 | 4.101 | 3.1003 |  | = Formulae | 2.8 | #NAME? | 19 | #NAME?
20 | 51.35 | 18.3393 | 5.045 | 2.909 |  |  | 2.9 | #NAME? | 20 | #NAME?
21 | 42.4 | 15.1429 | 4.976 | 2.7175 |  | = Reference Figures | 3 | #NAME? | 21 | #NAME?
22 | 35.01 | 12.5036 | 5.508 | 2.526 |  |  | 3.1 | #NAME? | 22 | #NAME?
23 | 28.95 | 10.3393 | 5.941 | 2.336 |  | = Manual Input | 3.2 | #NAME? | 23 | #NAME?
24 | 23.99 | 8.5679 | 6.573 | 2.148 |  |  | 3.3 | #NAME? | 24 | #NAME?
25 | 19.95 | 7.125 | 6.945 | 1.9636 | Doc. No: 200.XXX |  | 3.4 | #NAME? | 25 | #NAME?
26 | 16.66 | 5.95 | 7.261 | 1.7834 | Sheet 3 |  | 3.5 | #NAME? | 26 | #NAME?
27 | 13.98 | 4.9929 | 8.03 | 1.608 | DCRXXXX |  | 3.6 | #NAME? | 27 | #NAME?
28 | 11.8 | 4.2143 | 8.369 | 1.4385 | Rev 0 |  | 3.7 | #NAME? | 28 | #NAME?
29 | 10.02 | 3.5786 | 8.595 | 1.275 |  |  | 3.8 | #NAME? | 29 | #NAME?
30 | 8.568 | 3.06 | 9.943 | 1.1184 |  |  | 3.9 | #NAME? | 30 | #NAME?
31 | 7.38 | 2.6357 |  | 0.96915 |  |  | 4 | #NAME? | 31 | #NAME?
32 | 6.403 | 2.2868 |  | 0.82715 |  |  | 4.1 | #NAME? | 32 | #NAME?
33 | 5.594 | 1.9979 |  | 0.69208 |  |  | 4.2 | #NAME? | 33 | #NAME?
34 | 4.921 | 1.7575 |  | 0.56389 |  |  | 4.3 | #NAME? | 34 | #NAME?
35 | 4.356 | 1.5557 |  | 0.44193 |  |  | 4.4 | #NAME? | 35 | #NAME?
36 | 3.88 | 1.3857 |  | 0.32622 |  |  | 4.5 | #NAME? | 36 | #NAME?
37 | 3.476 | 1.2414 |  | 0.21626 |  |  | 4.6 | #NAME? | 37 | #NAME?
38 | 3.131 | 1.1182 |  | 0.11173 |  |  | 4.7 | #NAME? | 38 | #NAME?
39 | 2.834 | 1.0121 |  | 0.01207 |  |  | 4.8 | #NAME? | 39 | #NAME?
40 | 2.577 | 0.92036 |  | -0.082993 |  |  | 4.9 | #NAME? | 40 | #NAME?
41 | 2.354 | 0.84071 |  | -0.1735 |  |  | 5 | #NAME? | 41 | #NAME?
42 | 2.159 | 0.77107 |  | -0.25997 |  |  | 5.1 | #NAME? | 42 | #NAME?
43 | 1.988 | 0.71 |  | -0.34249 |  |  | 5.2 | #NAME? | 43 | #NAME?
44 | 1.836 | 0.65571 |  | -0.42203 |  |  | 5.3 | #NAME? | 44 | #NAME?
45 | 1.702 | 0.60786 |  | -0.49782 |  |  | 5.4 | #NAME? | 45 | #NAME?
46 | 1.581 | 0.56464 |  | -0.57156 |  |  | 5.5 | #NAME? | 46 | #NAME?
47 | 1.474 | 0.52643 |  | -0.64164 |  |  | 5.6 | #NAME? | 47 | #NAME?
48 | 1.377 | 0.49179 |  | -0.70971 |  |  | 5.7 | #NAME? | 48 | #NAME?
49 | 1.289 | 0.46036 |  | -0.77575 |  |  | 5.8 | #NAME? | 49 | #NAME?
50 | 1.209 | 0.43179 |  | -0.83983 |  |  | 5.9 | #NAME? | 50 | #NAME?
51 | 1.12 | 0.4 | MP Values | -0.91629 |  |  | 6 | #NAME? | 51 | #NAME?
52 | 1.064 | 0.38 | MP Values | -0.96758 |  |  | 6.1 | #NAME? | 52 | #NAME?
53 | 1.008 | 0.36 | MP Values | -1.0217 |  |  | 6.2 | #NAME? | 53 | #NAME?
54 | 0.952 | 0.34 | MP Values | -1.0788 |  |  | 6.3 | #NAME? | 54 | #NAME?
55 | 0.896 | 0.32 | MP Values | -1.1394 |  |  | 6.4 | #NAME? | 55 | #NAME?
56 | 0.84 | 0.3 | MP Values | -1.204 |  |  | 6.5 | #NAME? | 56 | #NAME?
57 | 0.812 | 0.29 | MP Values | -1.2379 |  |  | 6.6 | #NAME? | 57 | #NAME?
58 | 0.756 | 0.27 | MP Values | -1.3093 |  |  | 6.7 | #NAME? | 58 | #NAME?
59 | 0.7 | 0.25 | MP Values | -1.3863 |  |  | 6.8 | #NAME? | 59 | #NAME?
60 | 0.644 | 0.23 | MP Values | -1.4697 |  |  | 6.9 | #NAME? | 60 | #NAME?
61 | 0.588 | 0.21 | MP Values | -1.5606 |  |  | 7 | #NAME? | 61 | #NAME?
62 | 0.56 | 0.2 | MP Values | -1.6094 |  |  | 7.1 | #NAME? | 62 | #NAME?
63 | 0.532 | 0.19 | MP Values | -1.6607 |  |  | 7.2 | #NAME? | 63 | #NAME?
64 | 0.504 | 0.18 | MP Values | -1.7148 |  |  | 7.3 | #NAME? | 64 | #NAME?
65 | 0.448 | 0.16 | MP Values | -1.8326 |  |  | 7.4 | #NAME? | 65 | #NAME?
66 | 0.392 | 0.14 | MP Values | -1.9661 |  |  | 7.5 | #NAME? | 66 | #NAME?
67 | 0.364 | 0.13 | MP Values | -2.0402 |  |  | 7.6 | #NAME? | 67 | #NAME?
68 | 0.336 | 0.12 | MP Values | -2.1203 |  |  | 7.7 | #NAME? | 68 | #NAME?
69 | 0.308 | 0.11 | MP Values | -2.2073 |  |  | 7.8 | #NAME? | 69 | #NAME?
70 | 0.28 | 0.1 | MP Values | -2.3026 |  |  | 7.9 | #NAME? | 70 | #NAME?
71 | 0.252 | 0.09 | MP Values | -2.4079 |  |  | 8 | #NAME? | 71 | #NAME?
72 | 0.224 | 0.08 | MP Values | -2.5257 |  |  | 8.1 | #NAME? | 72 | #NAME?
73 | 0.196 | 0.07 | MP Values | -2.6593 |  |  | 8.2 | #NAME? | 73 | #NAME?
74 | 0.1876 | 0.067 | MP Values | -2.7031 |  |  | 8.3 | #NAME? | 74 | #NAME?
75 | 0.1792 | 0.064 | MP Values | -2.7489 |  |  | 8.4 | #NAME? | 75 | #NAME?
76 | 0.1708 | 0.061 | MP Values | -2.7969 |  |  | 8.5 | #NAME? | 76 | #NAME?
77 | 0.1624 | 0.058 | MP Values | -2.8473 |  |  | 8.6 | #NAME? | 77 | #NAME?
78 | 0.154 | 0.055 | MP Values | -2.9004 |  |  | 8.7 | #NAME? | 78 | #NAME?
79 | 0.1456 | 0.052 | MP Values | -2.9565 |  |  | 8.8 | #NAME? | 79 | #NAME?
80 | 0.1372 | 0.049 | MP Values | -3.0159 |  |  | 8.9 | #NAME? | 80 | #NAME?
81 | 0.1288 | 0.046 | MP Values | -3.0791 |  |  | 9 | #NAME? | 81 | #NAME?
82 | 0.1204 | 0.043 | MP Values | -3.1466 |  |  | 9.1 | #NAME? | 82 | #NAME?
83 | 0.112 | 0.04 | MP Values | -3.2189 |  |  | 9.2 | #NAME? | 83 | #NAME?
84 | 0.1036 | 0.037 | MP Values | -3.2968 |  |  | 9.3 | #NAME? | 84 | #NAME?
85 | 0.0952 | 0.034 | MP Values | -3.3814 |  |  | 9.4 | #NAME? | 85 | #NAME?
86 | 0.0868 | 0.031 | MP Values | -3.4738 |  |  | 9.5 | #NAME? | 86 | #NAME?
87 | 0.0784 | 0.028 | MP Values | -3.5756 |  |  | 9.6 | #NAME? | 87 | #NAME?
88 | 0.07 | 0.025 | MP Values | -3.6889 |  |  | 9.7 | #NAME? | 88 | #NAME?
89 | 0.0616 | 0.022 | MP Values | -3.8167 |  |  | 9.8 | #NAME? | 89 | #NAME?
90 | 0.056 | 0.02 | MP Values | -3.912 |  |  | 9.9 | #NAME? | 90 | #NAME?

## ICM Job Card 2022 (hidden)
[9 empty] | 202.027_02 |  | - | - | - | -
[9 empty] | DCR0678 | 3900 | C Jones | ISO4406:1999 | N | MK II
Serial Number |  |  |  | Product Code | [6 empty] | B Brown | NAS1638 | 0 | MK I
Customer | [10 empty] | A Hayward | AS4059E Table 2 | 1 | MK I (OLD)
Booking Date |  |  |  | Service No. | [6 empty] | A Wallace | AS4095E Table 1 | 2
Sales Order |  |  |  | Works Order | [6 empty] | A Mustoe | ISO11218 | 3
Box Kept? | [10 empty] | J Scarrott | AS4059F Table 2 | 4
Checks carried out by |  |  | - | [7 empty] | S Mosson | AS4095F Table 1 | 5
Calibration Screen |  | Initial | Rig - (oil flow) |  | Service carried out by: |  | - |  |  |  | A Buckley |  | 6
Photocurrent | (nA) |  |  |  | Calibration carried out by: |  | - |  |  |  | T Whittle |  | 7
RH % |  |  |  |  | Continuous Testing Options |  |  | Additional notes
Temp |  |  |  |  | Test continuously |  | FALSE
User Settings |  |  |  |  | Log continuously |  | FALSE
Format |  | - |  |  | Start testing automatically |  | FALSE
Simulated Test |  | FALSE |  |  | Stop testing when clean |  | FALSE
Low Flow Alarm Disabled |  | FALSE |  |  | Confirm target levels |  | FALSE
Alarm Mode |  | - |  |  | Before stopping (x) times?
Ignore initial tests
System Settings
Sensing Vol. Parameters |  | - |  |  | Firmware Updated
Test Reference |  |  |  |  | Y/N?
Accumulation Time (s)
Auto-repeat on? |  | FALSE |  |  | Pressure Tests
Interval (s) |  |  |  |  | 500b Static
W - Water Sensor |  | FALSE |  |  | 500b Pulse
R - Alarm Relays |  | FALSE
Contamination Code Target/Alarm Levels
| 4 | 6 | 14 |  | RH | Temp
Upper
Lower
Other Sizes
Electronic Changes |  | PCBR | PCBT | PCBK |  | Date Returned
Serial numbers
Note on F21: Tick box if you have updated the firmware

## Service Report (hidden)
Serial Number | 0
Service Number | 0
Customer | 0
Sales Order | 0
Product | 0
Faults Identified
Work Carried Out / Faulty Parts Replaced
Further Recommendations
Service Carried Out By | -

## UK Sterling service quote (hidden)
Service Quotation: |  |  | #REF!
Serial number of unit(s): |  |  | #REF!
Customer: |  | #REF!
Date: | #REF!
Required work:
Fault Found
0
0
0
0
0
0
Work Done
0
0
0
0
0
0
Cost details: |  |  | All figures are in GBP £
Stock Code |  | Description |  | Qty | Price | Discount % | Net Price
15.005 |  | ICM Recalibration |  | 1 | 178 |  | 178
0 |  | 0 |  | 0 | 0 |  | 0
0 |  | 0 |  | 0 | 0 |  | 0
0 |  | 0 |  | 0 |  |  | 0
0 |  | 0 |  | 0 |  |  | 0
0
0
0
0
0
0
0
[6 empty] | Total | 178
[6 empty] | VAT | 35.6
[6 empty] | GROSS | 213.6
Customer Authorisation |  |  | Please sign and return this authorisation to swillett@mpfiltri.co.uk or sales@mpfiltri.co.uk
We accept the above cost to be added to any current order lines, please advise if this isn't acceptable
We accept the above service quotation and authorise MP Filtri to proceed with the repairs
Signed………………………………………………………………………………………………………………. |  |  |  |  | Service quote number |  | #REF!
[5 empty] | Date of service quote |  | #REF!
Note on J30: Steven Willett: DO NOT ALTER THESE FIELDS

## Non UK Sterling service quote (hidden)
Service Quotation: |  |  | #REF!
Serial number of unit(s): |  |  | #REF!
Customer: |  | #REF!
Date: | #REF!
Required work:
0
0
0
0
0
Fault Found
Work Done
0
0
0
0
0
0
0
Cost details: |  |  | All figures are in GBP £
Stock Code |  | Description |  | Qty | Price |  | Discount % | Net Price
15.005 |  | ICM Recalibration |  | 1 | 178 |  |  | 178
0 |  | 0 |  | 0 | 0 |  |  | 0
0 |  | 0 |  | 0 | 0 |  |  | 0
0 |  | 0 |  | 0 | 0 |  |  | 0
0 |  | 0 |  | 0 | 0 |  |  | 0
|  | 0 |  | 0 | 0 |  |  | 0
0
0
0
0
0
0
[7 empty] | Total | 178
Customer Authorisation |  |  | Please sign and return this authorisation to swillett@mpfiltri.co.uk or sales@mpfiltri.co.uk
We accept the above service quotation and authorise MP Filtri to proceed with the repairs
Signed………………………………………………………………………………………………………………. | [5 empty] | Service quote number |  | #REF!
[6 empty] | Date of service quote |  | #REF!

## Euro Service Quote (hidden)
Service Quotation: |  |  | #REF!
Serial number of unit(s): |  |  | #REF!
Customer: |  | #REF!
Date: | #REF!
Required work:
0
0
0
0
0
Fault Found
Work Done
0
0
0
0
0
0
0
Cost details: |  |  | All figures are in GBP £
Stock Code |  | Description |  | Qty | Price |  | Discount % | Net Price
15.005 |  | ICM Recalibration |  | 1 | 200 |  |  | 200
0 |  | 0 |  | 0 | 0 |  |  | 0
0 |  | 0 |  | 0 | 0 |  |  | 0
0 |  | 0 |  | 0 | 0 |  |  | 0
0 |  | 0 |  | 0 | 0 |  |  | 0
|  |  |  | 0 |  |  |  | 0
0
0
0
0
0
0
[7 empty] | Total | 200
Customer Authorisation |  |  | Please sign and return this authorisation to swillett@mpfiltri.co.uk or sales@mpfiltri.co.uk
We accept the above service quotation and authorise MP Filtri to proceed with the repairs
Signed………………………………………………………………………………………………………………. | [5 empty] | Service quote number |  | #REF!
[6 empty] | Date of service quote |  | #REF!

## Calibration record
[7 empty] | = Locked cells |  |  | 202.027
Rev 1
= Manual Input
[6 empty] | CMP type | 11.1111
[6 empty] | Serial Number | 000000204 / 07093-41
|  | Units |  |  |  | Date | 08 12 25
| 61 | - | No Noise Threshold |  |  | Cert No. | 000000204 / 07093-41 | 08 12 25
| 11558.8984 | nA | Photocurrent |  |  | MTD Batch No. | 12798M
| 15.9952 | mg/L | Calculated final concentration |  |  | Test Rig No. | 367
[6 empty] | Calibrated by | J Scarrott
(Light Source, LS1)
| 4 | 6 | 14 | 21 | 25 | 38 | 50 | 70 | Particle Size Required (µm)
| 134 | 324 | 1886 | 4122 | 6566 | 15500 | 22008 | 32390 | Set Thresholds
Add screen shots from LPA View of calibration here, include, LED Power level, and W results (where applicable)
Test result | Target | 1 | % Deviation 1 | 2 | % Deviation 2 | 3 | % Deviation 3 | Average Value | % Deviation Ave.
4 | 3481812.2857 | 3502561 | 0.5959% | 3504210 | 0.6433% | 3538386 | 1.625% | 3515052.3333 | 0.9547%
6 | 1368160.8571 | 1384415 | 1.188% | 1382038 | 1.014% | 1399993 | 2.327% | 1388815.3333 | 1.51%
14 | 97342.2171 | 97621 | 0.2864% | 99179 | 1.887% | 99930 | 2.658% | 98910 | 1.611%
21 | 24221.3029 | 23939 | -1.166% | 24541 | 1.32% | 25089 | 3.582% | 24523 | 1.246%
25 | 11396.58 | 11133 | -2.313% | 11031 | -3.208% | 11724 | 2.873% | 11296 | -0.8825%
38 | 1788.6061 | 1551 | -13.28% | 1458 | -18.48% | 1947 | 8.856% | 1652 | -7.638%
50 | 690.6499 | 593 | -14.14% | 482 | -30.21% | 379 | -45.12% | 484.6667 | -29.82%
70 | 159.952 | 71 | -55.61% | 71 | -55.61% | 35 | -78.12% | 59 | -63.11%
Add screen shots from LPA View of verification tests here
Note on A22: 10% range of values for 16mg/L: 3,140,000 - 3,830,000
Note on A23: 10% range of values for 16mg/L: 1,232,000 - 1,505,000
Note on A24: 10% range of values for 16mg/L: 88,000 - 107,000
Note on A25: 10% range of values for 16mg/L: 22,000- 26,500
Note on A26: 10% range of values for 16mg/L: 10,300 - 12,500
Note on A27: 10% range of values for 16mg/L: 1610 - 1960

## Post calibration (hidden)
Restoring customer settings
User Settings
| Format | - | FALSE
FALSE | Simulated Test | 2 | FALSE
FALSE | Low Flow Alarm Disabled | 2 | FALSE
| Alarm Mode | - | FALSE
| Contamination codes restored? |  | FALSE
Continuous Testing Options
FALSE | Test continuously | 2 | FALSE
FALSE | Log continuously | 2 | FALSE
FALSE | Start testing automatically | 2 | FALSE
FALSE | Stop testing when clean | 2 | FALSE
FALSE | Confirm target levels | 2 | FALSE
| Before stopping (x) times? | 0 | FALSE
| Ignore initial tests | 0 | FALSE
System Settings
| Sensing Vol. Parameters | - | FALSE
| Test Reference | 0 | FALSE
| Accumulation Time (s) | 0 | FALSE
FALSE | Auto-repeat on? | 2 | FALSE
| Interval (s) | 0 | FALSE
FALSE | W - Water Sensor | 2 | FALSE
FALSE | R - Alarm Relays | 2 | FALSE
Admin and record-keeping
| Complete calibration record |  | FALSE
| Select appropriate cal cert |  | FALSE
| Print to pdf |  | FALSE
| Print 2 physical copies in colour |  | FALSE
| Update simplybook.me |  | FALSE
| Fill in Job Card |  | FALSE
| Update whiteboard on window |  | FALSE
Packing
| Turn off rig, switch to clean-up |  | FALSE
| Close oil taps and turn rig back on |  | FALSE
| Attach bottle sampler to lines |  | FALSE
| Blow air through the ICMs to clear |  | FALSE
| Add calibration sticker under label |  | FALSE
| Remove and take to logistics |  | FALSE
| Post-calibration completed by: |  | -
FALSE
Note on C9: Check the contamination code targets in User Settings against the screenshot you took on the Job Card
Note on C32: Every orange box must be filled in, check CMP number and Serial number, see notes on the sheet for advice for filling out
Note on C33: Choose correct company ie CC Jensen, Descase for RMF etc. For Descase Cal cert, check print layout before printing
Note on C34: Save in Service & Recal folder in format: SN#_DDMMYY_"CC"_Service number
Note on C36: Green if SO has been generated, purple if still awaiting SO
Note on C37: Pressure tests, and fill in the rig (oil flow) section from the calibration screenshot you took
Note on C38: Add the number of units you have calibrated to the tally
Note on C43: Do NOT switch valve from test to clean-up while the rig is running
Note on C45: Connect outlet of bottle sampler to inlet of first ICM in line on top and bottom row
Note on C46: If the calibration is finished, you can blow air straight back into drum of rig

## MP Calibration Certificate
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park South,
Evesham
England
WR11 1 LB
Tel: +44 (0) 1386 258500
www.mpfiltri.co.uk
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | ICMWMKRG12.0AZ2 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | 08 12 25 |  |  |  |  | CALIBRATION RIG NUMBER |  | 367
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 1 YEAR FROM FIRST USE (refer to ICM diagnostics)
CALIBRATED BY | [6 empty] | Position | Calibration Lab Technician
J Scarrott
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 202.027-02
Sheet 1
DCRXXXX
Page 1 of 1 | [8 empty] | Rev 0

## DESCASE CERT (hidden)
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | 11.1111 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | 08 12 25 |  |  |  |  | CALIBRATION RIG NUMBER |  | 367
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 1 YEAR FROM FIRST USE (refer to CMS diagnostics)
CALIBRATED BY | J Scarrott | [5 empty] | Position | Calibration Lab Technician
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 202.027
Page 1 of 1

## LPM Certificate (hidden)
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | 11.1111 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | [5 empty] | CALIBRATION RIG NUMBER |  | 367
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 1 YEAR FROM FIRST USE (refer to ICM diagnostics)
CALIBRATED BY | J Scarrott | [5 empty] | Position | Calibration Lab Technician
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 202.027
Page 1 of 1

## Solge Certificate (hidden)
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | 11.1111 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | [5 empty] | CALIBRATION RIG NUMBER |  | 367
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 1 YEAR FROM FIRST USE (refer to ICM diagnostics)
CALIBRATED BY | J Scarrott | [5 empty] | Position | Calibration Lab Technician
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 202.027
Page 1 of 1

## RMF Certificate (hidden)
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | 11.1111 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | [5 empty] | CALIBRATION RIG NUMBER |  | 367
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 1 YEAR FROM FIRST USE (refer to ICM diagnostics)
CALIBRATED BY | J Scarrott | [5 empty] | Position | Calibration Lab Technician
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 202.027
Page 1 of 1

## CCJ Certificate (hidden)
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | 11.1111 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | [5 empty] | CALIBRATION RIG NUMBER |  | 367
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 2 YEARS FROM FIRST USE (refer to OCM diagnostics)
CALIBRATED BY | J Scarrott | [5 empty] | Position | Calibration Lab Technician
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 202.027
Page 1 of 1

## Hanzaflex Certificate (hidden)
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | 11.1111 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | [5 empty] | CALIBRATION RIG NUMBER |  | 367
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 1 YEAR FROM FIRST USE (refer to ICM diagnostics)
CALIBRATED BY | J Scarrott | [5 empty] | Position | Calibration Lab Technician
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 202.027
Page 1 of 1

## ICM SUNHML Certificate (hidden)
MP Filtri UK Ltd
Bourton on the Water Industrial Park
Gloucestershire
England
GL54 2HQ
Tel: +44 (0) 1451 822522
www.mpfiltri.co.uk
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | 11.1111 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | [5 empty] | CALIBRATION RIG NUMBER |  | 367
[7 empty] | AMBIENT TEMPERATURE OF TEST AREA |  | 21C
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATED FOR SUNHML ON BEHALF OF THE ROYAL SAUDI AIR FORCE, SAUDI ARABIA
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 1 YEAR FROM FIRST USE (refer to ICM diagnostics)
CALIBRATED BY | J Scarrott | [5 empty] | Position | Calibration Lab Technician
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 200.XXX
Sheet 1
DCRXXXX
Page 1 of 1 | [8 empty] | Rev 0

## Unbranded Cert (hidden)
CERTIFICATE OF CALIBRATION
WE HEREBY CERTIFY THAT THE CONTAMINATION MONITORING PRODUCT DETAILED BELOW HAS BEEN FULLY CALIBRATED WITH ISO MEDIUM TEST DUST (MTD) BASED ON ISO 11171 ON TEST EQUIPMENT CERTIFIED TO ISO 11943 BY I.F.T.S {Institut de la Filtration et des Techniques Séparatives}.
We hereby confirm that the above mentioned measuring system was calibrated according to DIN ISO 9001, under the observation of a certified quality assurance system. The measuring installation used for calibration are regularly calibrated to national standards by I.F.T.S. Should no national standards exist, the measuring procedure corresponds with the technical regulations and norms valid at the time of the measurement
PRODUCT TYPE | 11.1111 | [5 empty] | MEDIUM TEST DUST BATCH NUMBER |  | 12798M
SERIAL NUMBER | 000000204 / 07093-41 | [5 empty] | DATE OF CALIBRATION |  | 08 12 25
CERTIFICATE NUMBER | 000000204 / 07093-41 | [5 empty] | CALIBRATION RIG NUMBER |  | 367
PARTICLE COUNT LIMITS AT 16mg/L CONCENTRATION
(Suspension of ISO MTD in MIL-H-5606)
No. particles reported per 100ml volume
| Particle Size |  | Calibration Pulse Height |  | Target Particle Count | Actual Particle Count | Deviation | % Deviation (+/- 10%)
| [microns (c)] |  | [mV] |  | [n] | [n] | [n] | [%]
| >4 |  | 134 |  | 3481812.2857 | 3515052.3333 | 33240.0476 | 0.9547%
| >6 |  | 324 |  | 1368160.8571 | 1388815.3333 | 20654.4762 | 1.51%
| >14 |  | 1886 |  | 97342.2171 | 98910 | 1567.7829 | 1.611%
| >21 |  | 4122 |  | 24221.3029 | 24523 | 301.6971 | 1.246%
| >25 |  | 6566 |  | 11396.58 | 11296 | -100.58 | -0.8825%
| >38 |  | 15500 |  | 1788.6061 | 1652 | -136.6061 | -7.638%
| >**50 |  | 22008 |  | 690.6499 | 484.6667 | - | -
| >**70 |  | 32390 |  | 159.952 | 59 | - | -
**These particle sizes are not yet certified by NIST
| LS1 No Noise Threshold | LS1 Threshold Sensitivity |  | Sensitivity Threshold Ratio
| [mV] | [mV] |  | [-]
| 61 | 73 |  | 2.1967
CALIBRATION RESULT | PASS | [5 empty] | PASS - VERIFIED
RE-CALIBRATION DUE | 1 YEAR FROM FIRST USE (refer to ICM diagnostics)
CALIBRATED BY | [6 empty] | Position | Calibration Lab Technician
J Scarrott
Newly purchased products are covered under warranty for 12 months from first use | [6 empty] | We hereby certify that this instrument meets standard performance criteria, providing calibration results within +/-10% of target concentration values from >4um to >38um, and exhibits a sensitivity threshold ratio of greater than 1.19 for each light source
Products being returned for re-calibration are verified for accuracy for a further 12 months. This does not extend the warranty.
Customers are responsible for managing the intervals between calibration's as defined within IS0 9001:2000
Doc No. 200.XXX
Sheet 1
DCRXXXX
Page 1 of 1 | [8 empty] | Rev 0