## BS110 Cal'd Thresholds
Cal'd thresholds | [10 empty] | Counts
S/N | Ref material | 4 | 6 | 10 | 14 | 21 | 25 | 38 | 50 | 70 | 4 | 6 | 10 | 14 | 21 | 25 | 38 | 50 | 70
63010058 | 35D | 96 | 234 | 644 | 1252 | 2520 | 3710 | 10198 | 15876 | 25071 | 1067280 | 452414 | 107290 | 34817 | 10344 | 4626 | 438 | 85 | 4
63010051 | 35D | 92 | 212 | 578 | 1110 | 2210 | 3204 | 10011 | 16174 | 24822 | 1080432 | 464413 | 108641 | 37354 | 10130 | 5439 | 540 | 83 | 23
63010055 | 35D | 82 | 190 | 530 | 1030 | 2076 | 3036 | 10144 | 15958 | 25096 | 1098065 | 468172 | 107450 | 36852 | 11173 | 5463 | 195 | 37 | 4
63010046 | 35D | 90 | 218 | 600 | 1178 | 2406 | 3528 | 9952 | 16091 | 24765 | 1095753 | 463914 | 107953 | 35860 | 10524 | 4913 | 292 | 71 | 11
63010063 | 35D | 86 | 196 | 546 | 1076 | 2230 | 3408 | 10233 | 16112 | 25018 | 1081379 | 467448 | 107530 | 36531 | 10988 | 5210 | 397 | 146 | 40
63010045 | 35D | 92 | 212 | 578 | 1116 | 2204 | 3170 | 9921 | 15849 | 24883 | 1072065 | 463979 | 106685 | 36282 | 10974 | 5229 | 153 | 5 | 0
63010048 | 35D | 80 | 184 | 514 | 1006 | 2032 | 2962 | 10165 | 16205 | 24750 | 1084234 | 465319 | 108256 | 37308 | 11293 | 5214 | 140 | 35 | 5
63010053 | 35D | 88 | 198 | 546 | 1060 | 2192 | 3290 | 10037 | 16014 | 25093 | 1047957 | 452696 | 103307 | 35294 | 10125 | 4785 | 251 | 54 | 13
63010060 | 35D | 68 | 152 | 428 | 876 | 1886 | 2856 | 10256 | 15907 | 24810 | 1073037 | 463405 | 107139 | 36590 | 10984 | 5215 | 245 | 28 | 5
63010059 | 35D | 90 | 204 | 552 | 1064 | 2092 | 3036 | 9879 | 16188 | 24977 | 1067797 | 465672 | 109250 | 37641 | 12062 | 5929 | 280 | 59 | 5
63010062 | 35D | 92 | 206 | 562 | 1092 | 2204 | 3178 | 10118 | 15862 | 24791 | 1026589 | 439066 | 100360 | 34463 | 10400 | 5071 | 291 | 52 | 5
63010057 | 28DR | 74 | 184 | 599 | 1020 | 2080 | 2938 | 10084 | 16139 | 25112 | 583527 | 244923 | 52820 | 18403 | 5520 | 2422 | 176 | 51 | 11
63010064 | 28DR | 84 | 198 | 562 | 1126 | 2374 | 3670 | 9944 | 15971 | 24856 | 587439 | 248984 | 55040 | 18220 | 5552 | 2417 | 197 | 67 | 22
63010052 | 28DR | 70 | 176 | 538 | 1074 | 2394 | 3666 | 10201 | 16053 | 25029 | 581699 | 246254 | 55816 | 18361 | 4886 | 2297 | 178 | 44 | 11
63010054 | 28DR | 86 | 204 | 572 | 1108 | 2382 | 3608 | 10029 | 16221 | 24748 | 576773 | 241361 | 54256 | 18343 | 5165 | 2195 | 196 | 11 | 0
63010061 | 28DR | 90 | 206 | 566 | 1106 | 2376 | 3696 | 10177 | 15895 | 25088 | 568042 | 233547 | 54256 | 19116 | 5478 | 2206 | 146 | 33 | 5
63010049 | 28DR | 80 | 190 | 542 | 1084 | 2606 | 3570 | 9907 | 16076 | 24844 | 575120 | 240372 | 53477 | 17226 | 4766 | 1968 | 124 | 34 | 0
63010047 | 28DR | 84 | 194 | 544 | 1102 | 2350 | 3700 | 10245 | 15924 | 24995 | 564137 | 233524 | 51730 | 17526 | 5242 | 2362 | 213 | 50 | 11
63010056 | 28DR | 82 | 194 | 552 | 1132 | 2460 | 3678 | 9988 | 16167 | 24783 | 588203 | 241359 | 54502 | 18240 | 5174 | 2277 | 213 | 22 | 0
63010050 | 28DR | 88 | 198 | 534 | 1022 | 2070 | 3086 | 10132 | 15873 | 25054 | 580080 | 247742 | 57630 | 20111 | 6461 | 3057 | 159 | 51 | 0
[8 empty] | 10072 | 16030 | 24934
Step 1 |  | Prepare PartiStan - 60s shake, 60s ultrasonic, 60s turn over/roll bottle, 60s ultrasonic
Step 2 |  | Decant into 4x bottles (Cleaned with heptane + dried)
Step 3 |  | Open LPA View, connect LPA3 via USB-NETTWarrior to laptop. Ctrl+N > LPA3 Calibration > Select LPA3 COM port
Step 4 |  | Select PartiStan varient (EG, Partistan 2806 Lot 35D.) Click APPLY Edit
Step 5 |  | Shake 1x decanted bottle of fluid for 60s. Remove lid & put bottle into bottle sampler. Fit Vacuum cap 60s Vacuum. If bubbles still appear at top, another 15-30s vacuum.
Step 6 |  | switch bottle sampler cap - set to unit to sample. Turn on unit. Let bottle sampler unit it reaches pressure and stops
Step 7 |  | Start test on LPA3 once Bottle sampler switches off.
Step 8 |  | Allow LPA3 test to run
Step 9 |  | Once LPA3 says EMPTYING, click 'AUTO ADJUST' button on LPA view. allow 30~s for adjustment to complete.
Step 9 |  | Manually set 38, 50, 70 micron levels on the LPA3 itself. Then on the same screen, click 'recalculate'
Step 10 |  | Once test is finished, repeat steps 4-8. Await test results to auto print to compare as a verification test.

## Serial Number 63010045
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23001 |  |  |  | BS2139
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010045 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10720.65 |  | 4 | 10809.4 | 92
6 | 4647.7 | 4639.79 |  | 6 | 4647.7 | 212
10 | 1074.3 | 1066.85 |  | 10 | 1074.3 | 578
14 | 365.5 | 362.82 |  | 14 | 365.5 | 1116
21 | 109.8 | 109.74 |  | 21 | 109.8 | 2204
25 | 52.1 | 52.29 |  | 25 | 52.1 | 3170
38 | 438 | 441 |  | 38 | 438 | 9921
50 | 264 | 290 |  | 50 | 264 | 1589
70 | 87 | 80 |  | 70 | 87 | 24883
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010046
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23002 |  |  |  | BS2130
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010046 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10957.53 |  | 4 | 10809.4 | 90
6 | 4647.7 | 4639.14 |  | 6 | 4647.7 | 218
10 | 1074.3 | 1079.53 |  | 10 | 1074.3 | 600
14 | 365.5 | 358.6 |  | 14 | 365.5 | 1178
21 | 109.8 | 105.24 |  | 21 | 109.8 | 2406
25 | 52.1 | 49.13 |  | 25 | 52.1 | 3528
38 | 450 | 461 |  | 38 | 450 | 9952
50 | 286 | 254 |  | 50 | 286 | 16091
70 | 82 | 70 |  | 70 | 82 | 24765
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010047
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23003 |  |  |  | BS2131
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010047 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5641.37 |  | 4 | 5822.6 | 84
6 | 2422.5 | 2335.24 |  | 6 | 2422.5 | 194
10 | 543.8 | 517.3 |  | 10 | 543.8 | 544
14 | 182.5 | 175.26 |  | 14 | 182.5 | 1102
21 | 51.7 | 52.42 |  | 21 | 51.7 | 2350
25 | 22.8 | 23.62 |  | 25 | 22.8 | 3700
38 | 457 | 448 |  | 38 | 457 | 10245
50 | 263 | 242 |  | 50 | 263 | 15924
70 | 79 | 81 |  | 70 | 79 | 24995
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010048
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23004 |  |  |  | BS2133
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010048 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10842.34 |  | 4 | 10809.4 | 80
6 | 4647.7 | 4653.19 |  | 6 | 4647.7 | 184
10 | 1074.3 | 1082.56 |  | 10 | 1074.3 | 514
14 | 365.5 | 373.08 |  | 14 | 365.5 | 1006
21 | 109.8 | 112.93 |  | 21 | 109.8 | 2032
25 | 52.1 | 52.14 |  | 25 | 52.1 | 2962
38 | 400 | 371 |  | 38 | 400 | 10165
50 | 132 | 132 |  | 50 | 132 | 16205
70 | 20 | 14 |  | 70 | 20 | 24750
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010049
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23005 |  |  |  | BS2144
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010049 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5751.2 |  | 4 | 5822.6 | 80
6 | 2422.5 | 2403.72 |  | 6 | 2422.5 | 190
10 | 543.8 | 534.77 |  | 10 | 543.8 | 542
14 | 182.5 | 172.26 |  | 14 | 182.5 | 1084
21 | 51.7 | 47.66 |  | 21 | 51.7 | 2606
25 | 22.8 | 19.68 |  | 25 | 22.8 | 3570
38 | 394 | 392 |  | 38 | 394 | 9907
50 | 79 | 89 |  | 50 | 79 | 16076
70 | 11 | 7.2 |  | 70 | 11 | 24844
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010050
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23006 |  |  |  | BS2205
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010050 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5800.8 |  | 4 | 5822.6 | 88
6 | 2422.5 | 2477.42 |  | 6 | 2422.5 | 198
10 | 543.8 | 576.3 |  | 10 | 543.8 | 534
14 | 182.5 | 201.11 |  | 14 | 182.5 | 1022
21 | 51.7 | 64.61 |  | 21 | 51.7 | 2070
25 | 22.8 | 30.57 |  | 25 | 22.8 | 3086
38 | 441 | 481 |  | 38 | 441 | 10132
50 | 102 | 96 |  | 50 | 102 | 15873
70 | 23 | 12 |  | 70 | 23 | 25054
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010051
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23007 |  |  |  | BS2206
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010051 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10804.32 |  | 4 | 10809.4 | 92
6 | 4647.7 | 4644.13 |  | 6 | 4647.7 | 212
10 | 1074.3 | 1086.41 |  | 10 | 1074.3 | 578
14 | 365.5 | 373.54 |  | 14 | 365.5 | 1110
21 | 109.8 | 101.3 |  | 21 | 109.8 | 2210
25 | 52.1 | 54.39 |  | 25 | 52.1 | 3204
38 | 441 | 490 |  | 38 | 441 | 10011
50 | 102 | 132 |  | 50 | 102 | 16174
70 | 23 | 31 |  | 70 | 23 | 24822
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010052
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23008 |  |  |  | BS2208
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010052 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5816.99 |  | 4 | 5822.6 | 70
6 | 2422.5 | 2462.54 |  | 6 | 2422.5 | 176
10 | 543.8 | 558.16 |  | 10 | 543.8 | 538
14 | 182.5 | 183.61 |  | 14 | 182.5 | 1074
21 | 51.7 | 48.86 |  | 21 | 51.7 | 2394
25 | 22.8 | 22.97 |  | 25 | 22.8 | 3666
38 | 374 | 370 |  | 38 | 374 | 10201
50 | 102 | 102 |  | 50 | 102 | 16053
70 | 34 | 10 |  | 70 | 34 | 25029
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010053
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23009 |  |  |  | BS2137
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010053 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10479.57 |  | 4 | 10809.4 | 88
6 | 4647.7 | 4526.96 |  | 6 | 4647.7 | 198
10 | 1074.3 | 1033.07 |  | 10 | 1074.3 | 546
14 | 365.5 | 352.94 |  | 14 | 365.5 | 1060
21 | 109.8 | 101.25 |  | 21 | 109.8 | 2192
25 | 52.1 | 47.85 |  | 25 | 52.1 | 3290
38 | 289 | 271 |  | 38 | 289 | 10037
50 | 52 | 48 |  | 50 | 52 | 16014
70 | 5 | 3 |  | 70 | 5 | 25093
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010054
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23010 |  |  |  | BS2207
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010054 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5767.73 |  | 4 | 5822.6 | 86
6 | 2422.5 | 2413.61 |  | 6 | 2422.5 | 204
10 | 543.8 | 542.56 |  | 10 | 543.8 | 572
14 | 182.5 | 183.43 |  | 14 | 182.5 | 1108
21 | 51.7 | 51.65 |  | 21 | 51.7 | 2382
25 | 22.8 | 21.95 |  | 25 | 22.8 | 3608
38 | 359 | 303 |  | 38 | 359 | 10029
50 | 68 | 71 |  | 50 | 68 | 16221
70 | 12 | 11 |  | 70 | 12 | 24748
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010055
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23011 |  |  |  | BS2142
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010055 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10980.65 |  | 4 | 10809.4 | 82
6 | 4647.7 | 4681.72 |  | 6 | 4647.7 | 190
10 | 1074.3 | 1074.5 |  | 10 | 1074.3 | 530
14 | 365.5 | 368.52 |  | 14 | 365.5 | 1030
21 | 109.8 | 111.73 |  | 21 | 109.8 | 2076
25 | 52.1 | 54.63 |  | 25 | 52.1 | 3036
38 | 441 | 481 |  | 38 | 441 | 10144
50 | 102 | 124 |  | 50 | 102 | 15958
70 | 2.3 | 4.1 |  | 70 | 2.3 | 25096
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010056
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23012 |  |  |  | BS2138
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010056 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5882.03 |  | 4 | 5822.6 | 82
6 | 2422.5 | 2413.59 |  | 6 | 2422.5 | 194
10 | 543.8 | 545.02 |  | 10 | 543.8 | 552
14 | 182.5 | 182.4 |  | 14 | 182.5 | 1132
21 | 51.7 | 51.74 |  | 21 | 51.7 | 2460
25 | 22.8 | 22.77 |  | 25 | 22.8 | 3678
38 | 184 | 201 |  | 38 | 184 | 9988
50 | 34.5 | 33.4 |  | 50 | 34.5 | 16167
70 | 4.02 | 2.56 |  | 70 | 4.02 | 24783
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010057
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23013 |  |  |  | BS2134
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010057 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan 28DR
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5835.27 |  | 4 | 5822.6 | 74
6 | 2422.5 | 2449.23 |  | 6 | 2422.5 | 184
10 | 543.8 | 528.5 |  | 10 | 543.8 | 599
14 | 182.5 | 184.03 |  | 14 | 182.5 | 1020
21 | 51.7 | 55.2 |  | 21 | 51.7 | 2080
25 | 22.8 | 24.22 |  | 25 | 22.8 | 2938
38 | 374 | 401 |  | 38 | 374 | 10084
50 | 102 | 111 |  | 50 | 102 | 16139
70 | 34 | 36 |  | 70 | 34 | 25112
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010058
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23014 |  |  |  | BS2143
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010058 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10672.8 |  | 4 | 10809.4 | 96
6 | 4647.7 | 4524.14 |  | 6 | 4647.7 | 234
10 | 1074.3 | 1072.9 |  | 10 | 1074.3 | 644
14 | 365.5 | 348.17 |  | 14 | 365.5 | 1252
21 | 109.8 | 103.44 |  | 21 | 109.8 | 2520
25 | 52.1 | 46.26 |  | 25 | 52.1 | 3710
38 | 374 | 340 |  | 38 | 374 | 10198
50 | 102 | 121 |  | 50 | 102 | 15876
70 | 34 | 42 |  | 70 | 34 | 25071
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010059
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23015 |  |  |  | BS2140
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010059 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10677.97 |  | 4 | 10809.4 | 90
6 | 4647.7 | 4656.72 |  | 6 | 4647.7 | 204
10 | 1074.3 | 1092.5 |  | 10 | 1074.3 | 552
14 | 365.5 | 376.41 |  | 14 | 365.5 | 1064
21 | 109.8 | 120.62 |  | 21 | 109.8 | 2092
25 | 52.1 | 59.29 |  | 25 | 52.1 | 3036
38 | 394 | 397 |  | 38 | 394 | 9879
50 | 79 | 89 |  | 50 | 79 | 16188
70 | 11 | 25 |  | 70 | 11 | 24977
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010060
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23016 |  |  |  | BS2141
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010060 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10730.37 |  | 4 | 10809.4 | 68
6 | 4647.7 | 4634.05 |  | 6 | 4647.7 | 152
10 | 1074.3 | 1071.39 |  | 10 | 1074.3 | 428
14 | 365.5 | 365.9 |  | 14 | 365.5 | 876
21 | 109.8 | 109.84 |  | 21 | 109.8 | 1886
25 | 52.1 | 52.15 |  | 25 | 52.1 | 2856
38 | 480 | 405 |  | 38 | 480 | 10256
50 | 88 | 68 |  | 50 | 88 | 15907
70 | 18 | 12 |  | 70 | 18 | 24810
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010061
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23017 |  |  |  | BS2136
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-11
| Serial Number | 63010061 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan (28DR)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5680.42 |  | 4 | 5822.6 | 90
6 | 2422.5 | 2335.47 |  | 6 | 2422.5 | 206
10 | 543.8 | 542.56 |  | 10 | 543.8 | 566
14 | 182.5 | 191.16 |  | 14 | 182.5 | 1106
21 | 51.7 | 54.78 |  | 21 | 51.7 | 2376
25 | 22.8 | 22.06 |  | 25 | 22.8 | 3696
38 | 460 | 492 |  | 38 | 460 | 10177
50 | 88 | 64 |  | 50 | 88 | 15895
70 | 18 | 5 |  | 70 | 18 | 25088
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010062
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23018 |  |  |  | BS2135
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010062 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10265.89 |  | 4 | 10809.4 | 92
6 | 4647.7 | 4390.66 |  | 6 | 4647.7 | 206
10 | 1074.3 | 1003.6 |  | 10 | 1074.3 | 562
14 | 365.5 | 344.63 |  | 14 | 365.5 | 1092
21 | 109.8 | 104 |  | 21 | 109.8 | 2204
25 | 52.1 | 50.71 |  | 25 | 52.1 | 3178
38 | 184 | 143 |  | 38 | 184 | 10118
50 | 35 | 23 |  | 50 | 35 | 15862
70 | 4 | 9 |  | 70 | 4 | 24791
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010063
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23019 |  |  |  | BS2209
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010063 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | S250602008 (35D)
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 10809.4 | 10813.79 |  | 4 | 10809.4 | 86
6 | 4647.7 | 4674.48 |  | 6 | 4647.7 | 196
10 | 1074.3 | 1075.3 |  | 10 | 1074.3 | 546
14 | 365.5 | 365.31 |  | 14 | 365.5 | 1076
21 | 109.8 | 109.88 |  | 21 | 109.8 | 2230
25 | 52.1 | 52.1 |  | 25 | 52.1 | 3408
38 | 433 | 401 |  | 38 | 433 | 10233
50 | 83 | 68 |  | 50 | 83 | 16112
70 | 14 | 17 |  | 70 | 14 | 25018
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.

## Serial Number 63010064
MP Filtri UK Ltd
Keep House, Conference Way, Vale Park
Evesham
England
WR11 1LB
www.mpfiltri.com
NAVAIR Kit Serial Number Tracking
NAVAIR Particle Counter (S/N) |  |  |  | Bottle Sampler (S/N)
D23020 |  |  |  | BS2132
Table 4 — APC Calibration summary
APC Sensor | Model | LPA 3 NAVAIR |  | Date |  | 2026-06-10
| Serial Number | 63010064 |  | Operator |  | KAP
| Model | LPA 3 PRIMARY |  | Calibration type (primary or secondary) |  | SECONDARY
| Serial Number | 11171-00009 |  | Calibration sample identification numbers |  | PartiStan 28DR
Noise level (Annex A) [mV] |  | 54 |  | Coincidence error limit (Annex B) |  | 30000 | particles/mL
Sample volume (Annex A) [mL] |  | 40 |  | Working flow rate (Annex C) |  | 30 | mL/min
Cv, vol (Annex A) [%] |  | 1.2% |  | Flow rate range (Annex C) |  | 30 | mL/min
Resolution (Annex D)
S_R | 0.5 |  | µm | R_R | 2.7 |  | %
S_L | 0.6 |  | µm | R_L | 8.3 |  | %
d | 10.2 |  | µm | R | 8.3 |  | %
Verification of Particle Counting Accuracy (Annex E) |  |  |  | Particle Size Calibration Summary (Clause 6)
Particle Size | Expected particle concentration | Observed particle concentration |  | Particle Size | Particle Concentration | Threshold Setting
µm(c) | Particles/mL | Particles/mL |  | µm(c) | Particles/mL | mV
4 | 5822.6 | 5874.39 |  | 4 | 5822.6 | 84
6 | 2422.5 | 2489.84 |  | 6 | 2422.5 | 198
10 | 543.8 | 550.4 |  | 10 | 543.8 | 562
14 | 182.5 | 182.2 |  | 14 | 182.5 | 1126
21 | 51.7 | 55.52 |  | 21 | 51.7 | 2374
25 | 22.8 | 24.17 |  | 25 | 22.8 | 3670
38 | 184 | 201 |  | 38 | 184 | 9944
50 | 35 | 54 |  | 50 | 35 | 15971
70 | 4 | 3 |  | 70 | 4 | 24856
Method of calibration and traceability: This device was calibrated using PartiStan 2806 Secondary reference fluid, Sizes and particle concentrations determined on a primary calibrated instrument using NIST SRM 2806d. Particle sizes 38, 50, 70 microns calibrated using NIST traceable glass spheres in accordance with clause 6.