## Annex A
Annex A Preliminary Check
Analysis using Bottle Samples and 25% coincidence RM8632a | [11 empty] | Concentrate Dilution Calculator |  |  |  |  | Concentrate Manufacturing Calculator
Can perform Annex C at same time using 5x bottles |  |  |  | Volumetric Sampling Accuracy |  |  | Volume Actually Measured |  |  |  |  | Concentrate |  | 100 | mg/l |  | Target Final Fluid Volume ml |  | 250
[12 empty] | Final fluid volume |  | 100 | ml |  | Target Concentration mg/l |  | 1000
APC Model |  |  | REF CTR |  |  |  | Prepare 5 bottles of 10um PSL spheres of concentration 2000/ml |  |  |  |  | Target concentration |  | 1 | mg/l |  | mg required |  | 250
APC S/N |  |  | 11171-000001 |  |  |  | Obtain counts using 10ml flush + 5x10ml portions
Sensor Model |  |  | REF CTR |  |  |  | Import all portion data into following sheet (A - B1) |  |  |  |  | Concentrate to add |  | 1 | ml
Sensor S/N |  |  | 11171-000001 |  |  |  | Adjust A - B1 to A-B5 sheets deleting outliers as per 6.6 |  |  |  |  | Diluent to add |  | 99 | ml
Date |  |  | 2026-06-23 |  |  |  | Enter expected concentration in cell below i.e. 2000
Operator |  |  | JMD |  |  |  | Can also export histogram data for one bottle and perform Annex D
Threshold Noise | [11 empty] | Alt Concentrate Dilution Calculator
Photocurrent |  |  | 10443 | [8 empty] | Concentrate mg/l |  | 100
4um(c) Coincidence Limit /ml |  |  | 30000 | [8 empty] | Diluent volume ml |  | 500
A.1 RM8632 4um count/mg/l |  |  | 6300 | [8 empty] | Target concentration mg/l |  | 6.5
RM8632x mg/l at 25% coincidence limit |  |  | 1.1905 | [8 empty] | Concentrate to add |  | 32.5
[12 empty] | Added |  | 3.27
Concentrate mg/l |  |  | 100 | [8 empty] | Final Concentration |  | 0.64975
Wanted fluid volume ml |  |  | 100
Target quantity of dust in dilution |  |  | 0.11905
Concentrate to add ml |  |  | 1.1905
Clean fluid to add |  |  | 98.8095 | Date | 2026-06-23 | [6 empty] | Threshold factor |  | 4.4583
|  |  |  | Operator | JMD | [8 empty] | 160
|  | /ml | Deviation | Particle Counter | 11171-000001 |  | Date | [6 empty] | 713.3357
| Portion 1 | 8793.1 | 0.8182% |  |  |  | Operator |  | JMD
| Portion 2 | 8667.1 | -0.6265% | 100ml weight | 85.4 |  | Particle Counter
| Portion 3 | 8677.9 | -0.5027% | Flush | 20 | ml
| Portion 4 | 8736.8 | 0.1727% | Sample | 50 | ml | AVERAGE 10um count/ml
| Portion 5 | 8733.8 | 0.1383% | Total Fluid Programmed | 70 | ml | Expected 10um count/ml |  | 2000
|  |  |  | Final Measured Weight | 25.88 | g | Capture Fraction |  | 0
| Mean | 8721.74 |  | Fluid Weight Taken | 59.52
| Sample StdDev | 45.5455 |  | ml taken | 69.6956 |  | Set "Beam Flow Correction" |  | -1000
|  |  |  | Error | -0.4349%
[9 empty] | COAST | [6 empty] | BRAKING
[9 empty] | MODE | [6 empty] | MODE
|  |  |  | 100ml weight | 85.11 |  | Density | Sample | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
|  |  |  | Flush | 20 | ml | 85.35 | Weight delta | 7.14 | 7.08 | 7.08 | 18.53 | 8.67 | 8.68 | 17.33 | 7 | 6.84 | 6.79 | 6.76 | 6.82
| CVVOL | 0.5222% | <3% = PASS | Sample | 50 | ml |  | Volume | 8.3656 | 8.2953 | 8.2953 | 21.7106 | 10.1582 | 10.1699 | 20.3046 | 8.2015 | 8.0141 | 7.9555 | 7.9203 | 7.9906 | 0 | 0
|  |  |  | Total Fluid Programmed | 70 | ml |  | Programmed | 10 | 10 | 10 | 20 | 10 | 10 | 20 | 10 | 10 | 10 | 10 | 10 | 10 | 10
|  |  |  | Final Measured Weight | 26.33 | g |  | Error ml | -1.6344 | -1.7047 | -1.7047 | 1.7106 | 0.15817 | 0.16989 | 0.30463 | -1.7985 | -1.9859 | -2.0445 | -2.0797 | -2.0094 | -10 | -10
|  |  |  | Fluid Weight Taken | 58.78
|  |  |  | ml taken | 69.0636
|  |  |  | Error | -1.338%
|  |  |  | 100ml weight | 86.24
|  |  |  | Flush | 20 | ml
|  |  |  | Sample | 50 | ml
|  |  |  | Total Fluid Programmed | 70 | ml
|  |  |  | Final Measured Weight | 27.6 | g
|  |  |  | Fluid Weight Taken | 58.64
|  |  |  | ml taken | 67.9963 |  |  |  | MODE 1: BRAKING
[9 empty] | MODE | Suction mode repeatability - 0ml flush, 10ml sample volume repeated tests
|  |  |  | Error | -2.862% |  | Density |  | MODE 1: 85.24
[7 empty] | Sample |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[7 empty] | Weight delta g |  | 6.87 | 6.85 | 6.79 | 6.76 | 6.75 | 6.79 | 6.71 | 6.74 | 6.74 | 6.7
|  |  |  | 100ml weight | 85.35 |  | Volume ml |  | 8.0596 | 8.0361 | 7.9657 | 7.9305 | 7.9188 | 7.9657 | 7.8719 | 7.9071 | 7.9071 | 7.8602
|  |  |  | Flush | 0 | ml | Programmed ml |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | STDDEV
|  |  |  | Sample | 10 | ml | Error ml |  | -1.9404 | -1.9639 | -2.0343 | -2.0695 | -2.0812 | -2.0343 | -2.1281 | -2.0929 | -2.0929 | -2.1398 | 0.065436 | ml
|  |  |  | Total Fluid Programmed | 10 | ml
|  |  |  | Final Measured Weight | 78.21 | g
|  |  |  | Fluid Weight Taken | 7.14
|  |  |  | ml taken | 8.3656
|  |  |  | Error | -1.6344 | ml
|  |  |  | Error | -16.34%
|  |  |  | 100ml weight | 85.35
|  |  |  | Flush | 20 | ml
|  |  |  | Sample | 50 | ml
|  |  |  | Total Fluid Programmed | 70 | ml
|  |  |  | Final Measured Weight | 26.33 | g
|  |  |  | Fluid Weight Taken | 59.02
|  |  |  | ml taken | 69.1506
|  |  |  | Error | -1.213%
[7 empty] | Density | g/100ml | 85.24 1: 85.68
|  |  |  | 100ml weight |  |  | Sample |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
|  |  |  | Flush | 20 | ml | Weight delta |  | 6.78 | 15.32 | 23.88 | 32.52 | 41.13 | 49.47 | 58.23 | 66.78 | 75.22 | 83.75
|  |  |  | Sample | 50 | ml | Volume |  | 7.9132 | 17.8805 | 27.8711 | 37.9552 | 48.0042 | 57.7381 | 67.9622 | 77.9412 | 87.7918 | 97.7474
|  |  |  | Total Fluid Programmed | 70 | ml | Programmed ml |  | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | STDDEV | Average
|  |  |  | Final Measured Weight | 27.6 | g | Error ml |  | -2.0868 | -2.1195 | -2.1289 | -2.0448 | -1.9958 | -2.2619 | -2.0378 | -2.0588 | -2.2082 | -2.2526 | 0.093241 | -2.1195
|  |  |  | Fluid Weight Taken | -27.6
|  |  |  | ml taken | #DIV/0!
|  |  |  | Error | #DIV/0!
[7 empty] | Density | g/100ml | 10: 85.68
[7 empty] | Sample |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[7 empty] | Weight delta |  | 8.24 | 8.15 | 8.21 | 8.17 | 8.16 | 8.11
[7 empty] | Volume |  | 9.6172 | 9.5121 | 9.5822 | 9.5355 | 9.5238 | 9.4655
[7 empty] | Programmed ml |  | 10 | 10 | 10 | 10 | 10 | 10 |  |  |  |  | STDDEV | Average
[7 empty] | Error ml |  | -0.38282 | -0.48786 | -0.41783 | -0.46452 | -0.47619 | -0.53455 |  |  |  |  | 0.05357 | -0.46063
Applied further 0.46ml correction
[7 empty] | Density | g/100ml | 85.68 1: 85.68
[7 empty] | Sample |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[7 empty] | Weight delta |  | 8.5 | 17.13 | 25.65 | 34.26 | 42.79 | 51.38 | 59.91 | 68.47 | 77 | 85.59
[7 empty] | Volume |  | 9.9206 | 19.993 | 29.937 | 39.986 | 49.9416 | 59.9673 | 69.923 | 79.9136 | 89.8693 | 99.895
[7 empty] | Programmed ml |  | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | STDDEV | Average
[7 empty] | Error ml |  | -0.079365 | -0.0070028 | -0.063025 | -0.014006 | -0.058357 | -0.03268 | -0.077031 | -0.086368 | -0.13072 | -0.10504 | 0.039137 | -0.065359
- Tare
- Weigh empty bottle
- Use measureing cylinder to add 100ml of fluid
- weigh full bottle
- Ensure particle counter lines and internal systems filled with
fluid from previousl cycling / flushing
- perform bottle sample test
- observe waste hose outlet ensure not dripping during sample
- weigh bottle after test
- check Error < 3%
Measure weight decrease for various small set sample volumes to identify offset.
20ml/min = 0.333ml/s. 0.3s deadtime = 0.100ml.
Pump wind-up will likely contribute more. 1 rev ~