## Sheet1
LPA3/BS500 counts with sample volume
| counts/ml | counts/100ml | 25ml |  |  |  | 25ml |  |  |  | 25ml |  |  |  | 50ml |  |  |  | 100ml
ISO CODE | 21/19/16 | 21/19/16 | 21/20/16 | % | -1/2 ISO CODE | +1/2 ISO CODE | 21/20/16 | % | -1/2 ISO CODE | +1/2 ISO CODE | 21/20/16 | % | -1/2 ISO CODE | +1/2 ISO CODE | 21/20/16 | % |  |  | 21/20/16 | % | Mean count | Mean %
4um | 10915.4 | 1091540 | 1484303 | 35.9825 | 984303 | 1984303 | 1546801 | 41.7081 | 1046801 | 2046801 | 1564898 | 43.3661 | 1064898 | 2064898 | 1535658 | 40.6873 | 1035658 | 2035658 | 1503539 | 37.7447 | 1527039.8 | 39.8977
6um | 4749.6 | 474960 | 633503 | 33.3803 | 508503 | 758503 | 656797 | 38.2847 | 531797 | 781797 | 677753 | 42.6969 | 552753 | 802753 | 657222 | 38.3742 | 532222 | 782222 | 634250 | 33.5376 | 651905 | 37.2547
14um | 378.9 | 37890 | 40226 | 6.1652 | 22226 | 58226 | 42324 | 11.7023 | 24324 | 60324 | 42138 | 11.2114 | 24138 | 60138 | 44472 | 17.3713 | 26472 | 62472 | 35042 | -7.5165 | 40840.4 | 7.7868
21um | 99.2 | 9920 | 10079 | 1.6028 |  |  | 10415 | 4.9899 |  |  | 10929 | 10.1714 |  |  | 10712 | 7.9839 |  |  | 7270 | -26.7137 | 9881 | -0.39315
25um | 51.1 | 5110 | 4677 | -8.4736 |  |  | 3683 | -27.9256 |  |  | 5250 | 2.7397 |  |  | 4697 | -8.0822 |  |  | 2624 | -48.6497 | 4186.2 | -18.0783
38um | n/a | n/a | 544 | #VALUE! |  |  | 452 | #VALUE! |  |  | 503 | #VALUE! |  |  | 316 | #VALUE! |  |  | 112 | #VALUE!
50um | n/a | n/a | 201 | #VALUE! |  |  | 301 | #VALUE! |  |  | 503 | #VALUE! |  |  | 100 | #VALUE! |  |  | 24 | #VALUE!
70um | n/a | n/a | 0 | #VALUE! |  |  | 0 | #VALUE! |  |  | 0 | #VALUE! |  |  | 0 | #VALUE! |  |  | 0 | #VALUE!
Conclusion: Our MTD has less particles at 4um causing us to set a threshold such that we are counting a smaller size. This elevates the 4um count by 40% when testing Partistan fluid.
We can correct for this by correcting the MTD distribution we are using.
There are 40% more SRM2806 particles at ~3.4um CF 4um
So if we are counting 3.4um particles this is a circular equivalent area smaller by a factor 0.7225.
There are about 30% fewer 4um particles than 3.4um so we should reduce our MTD distribution by this amount at 4um i.e. multiply by 0.714
Similarly for 6um we are counting at ~5.2um
We should reduce the 6um distribution value by 27% i.e. a factor 0.73.