# LPA3-NAVAIR Online Calibration Workflow

2025-05-13

TODO:

Firmware Release

Ensure PLA5 USB programmed to LPA3

## FOR EACH UNIT

- Update firmware to

- Check appears as "LPA3" on connect dialog

System Settings

- Change "Product" to LPA3-NAVAIR. Unit will reset.

- Initialise Product Settings

- Set serial number according to the one on the Block

Leave flushing with clean fluid for at least 20 minutes. Leave with fluid in unit for subsequent calibration.

# Calibration

Setup

Use dialog to dose rig to 5mg/l MTD, "corrected" distribution tickbox on.

Distribution grid will populate and a line written to log indicating current target

# For Each Unit

Perform no-noise check a few times with clean fluid in block but test rig isolated (do this before calibration fluid through the unit.

Record Photocurrent and lowest noise value against serial number.

Connect to dosed test rig

Flush 100ml into waste

Flush for 1 minute

CAL START

38+μm thresholds will not be reliable especially at 5mg/l

These need to be manually loaded from reference values.

As of 2025-05-13, thresholds for LPA3-NAVAIR are:

| 38μm | 9938 |
| --- | --- |
| 50μm(c) | 15300 |
| 70μm(c) | 25000 |

(CF automatically calibrated values)